I.  General information:

1.  Mission name:  SENEX 2013

2.  Instrument name: CO2/CH4

3.  What is measured: CO2 and CH4

4.  Short description of measurement technique:

wavelength scanned cavity ring down spectroscopy

5. Contact information for all personnel going to the field with this instrument:

(for multiple investigators, please list the PI or primary contact person first)

| Name | Email | Office phone | Cell phone |
| --- | --- | --- | --- |
| Jeff Peischl | jeff.peischl@noaa.gov | 303-497-4849 | 303-246-6651 |
| Tom Ryerson | thomas.b.ryerson@noaa.gov | 303-497-7531 |  |

II.  Specific information:

1. Total installed weight:	256.9 lbs

(rack, gas cylinders, hoses, cabling, pumps, inlets, permeation tubes, etc.)

| Rack weight and balance info | Allowed | Actual |
| --- | --- | --- |
| Weight, lbs.: | 350 | 256.9 |
| Overturning moment, in-lbs.: | 9800 | 5608 |

Pod weight and CG:

NOTE: Please also provide weight-and-balance information for all installed equipment.  Templates for standard electronics racks are available for download here.  PIs with non-standard installations will need to provide relevant information in a similar format.

2.  Individual subassembly info (weights should sum to total listed above)

| Component name | Location name and flight station | Weight, lbs |
| --- | --- | --- |
| Instrument | Sta5 | 60.0 |
| Laptop/ethernet | Sta5 | 5.5 |
| Gas deck, 1 cyl. | Sta5 | 49.5 |
| Gas deck, 2 cyl. | Sta5 | 69.5 |
| pump | Sta5 | 23.1 |
| UPS Dry ice dewar Additional cabling Additional plumbing | Sta5 Sta5 Sta5 Sta5 | 31.3 3.0 5.0 10.0 |

3.  Component power consumption in Amps;

| Component name | Location name | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |
| --- | --- | --- | --- | --- | --- | --- |
| Instrument | Sta5 |  |  | 3.0 A |  |  |
| Gas deck | Sta5 |  |  | 0.5 A |  |  |
| Laptop | Sta5 |  |  | 1.5 A |  |  |
| pump | Sta5 |  |  |  | 4.0 A |  |
|  | Totals: |  |  | 5 A | 4 A |  |
|  |  | 400 Hz 3Ø | 400 Hz 1Ø | 60 Hz | 28VDC | 28VDC WOW |

4.  Inlet and exhaust information:

Inlet/exhaust line diagram in Appendix B

| Inlet/exhaust name | Location name and flight station | Hole size through hull, inches |
| --- | --- | --- |
| 1. CO2/CH4 | Sta5 | 3/8” inlet |

5.  Source of flow (name and location of pump or venturi)

| Pump name | Location name and flight station |
| --- | --- |
| 1. CO2/CH4 pump | Sta5 |

6.  Installed hazardous materials or equipment:

(only for items installed in the aircraft for use during flight)

A.  Lasers:

Type: TBD

Class: TBD

Wavelength: 1.60 – 1.65um

Output power: < 25 mW

Eye-safe?  yes

Beam fully contained within instrument during normal operation? yes

For non-eye-safe lasers, please attach a description of safety measures taken (safety interlocks, beam fully enclosed within instrument, etc.) and a procedure for safe instrument operation during testing and laser alignment.  Please contact the AIC for an example of laser safety documentation from TexAQS 2006.

B.  RF transmitters: N/A

C.  Radioactive materials: N/A

D.  Compressed gases: (1 ft3 = 28.32 liters; cabin volume = 4260 ft3 = 1.21 x 105 liters)

| Cylinder number: | 1 | 2 | 3 | 4 | 5 |
| --- | --- | --- | --- | --- | --- |
| Gas description | Compressed air | Compressed air | Compressed air |  |  |
| Mixing ratio | ~360 ppm CO2, ~1700 ppb CH4 | ~390 ppm CO2, ~1900 ppb CH4 | ~490 ppm CO2, ~2400 ppb CH4 |  |  |
| Cylinder size (ft3) | 30 | 30 | 30 |  |  |
| Max pressure (psig) | 2200 | 2200 | 2200 |  |  |
| # installed on aircraft | 1 | 1 | 1 |  |  |
| Location on aircraft | Sta5 | Sta5 | Sta5 |  |  |
| Service frequency | ~3 wks. | ~3 wks. | ~3 wks. |  |  |
| toxic/flammable gases: |  |  |  |  |  |
| In containment vessel? | no | no | no |  |  |
| Gas alarm provided? | no | no | no |  |  |
| MR if vented to cabin, ppmv | same | same | same |  |  |
| OSHA 8-hr PEL, ppmv | CO2: 5000 CH4: simple asphyxi-ant | CO2: 5000 CH4: simple asphyxi-ant | CO2: 5000 CH4: simple asphyxi-ant |  |  |
| 30-min IDLH, ppmv | CO2: 50000 CH4: simple asphyxi-ant | CO2: 50000 CH4: simple asphyxi-ant | CO2: 50000 CH4: simple asphyxi-ant |  |  |

E.  Chemicals (solids and liquids):

| Chemical number: |  |
| --- | --- |
| Concentration |  |
| Amount |  |
| Container description |  |
| Purpose |  |
| Solution pH |  |
| Spill kit provided? |  |

F.  Cryogens: N/A

G.  UPS and battery installation:

Location: Sta5

Description:  (Manufacturer, model no., power) Tripp Lite Smart1000, 1000 VA

Battery type: sealed lead-acid

Has an adjustable input voltage tolerance? (highly recommended)!

H.  Motors/Pumps

Description: Diaphragm pump, KNF PM24731-920

Motor current draw (28V, 7 A startup, ~4 A running)

Thermal interlock enabled? No

I. Operator seat requests 0

Test flights: 1

Transit flights: 0

Science flights: only when instrument issues are evident

7.  Data and plumbing drops

Network (Cat. 5/6 ethernet) drops requested: 1

Serial drops requested: 0

IRIG-B drops (BNC coax connector) requested: no

Vacuum/exhaust/ emergency dump lines:

| Name | Flow rate, slpm | Line pressure, Torr | Pump type | Trace gas concentration(s) |
| --- | --- | --- | --- | --- |
| 1. CO2/CH4 exhaust | ~1.25 | ambient | diaphragm pump | ambient |
| 2. |  |  |  |  |
| 3. |  |  |  |  |
| 4. |  |  |  |  |

Ground gas service lines (number, location, type of service): none

Other gas lines (number, location, type of service): none

Will you be sending data to the AOC data station?  If so, please provide the following information: no

| Parameter name | Voltage range | Unit conversion |
| --- | --- | --- |
| 1. |  |  |
| 2. |  |  |
| 3. |  |  |

8.  Aircraft access

a.  flight days:

Pre-flight time requested at aircraft (hours): 3h

Routine pre-flight ground support required? none

(stands, ladders, forklifts, covers, external equipment, etc.)

Routine post-flight time requested at aircraft (hours): 30 min

Routine post-flight ground support required? no

(stands, ladders, forklifts, covers, external equipment, etc.)

b.  non-flight days:

Routine external access to inlets or zenith mounts required? no

(please describe location, how often, for how long, type of ground support equipment needed, weather constraints, etc.)

Please note there is zero access and zero power to the aircraft (including pods) on hard-down days.  These occur at least once every seven calendar days while in the field.

9.  Aircraft maneuvers

Briefly describe in-flight calibration frequency, duration, altitudes desired:

calibration every hour or so for approximately 6 minutes at level altitude

Briefly describe instrument sensitivity to flight conditions:

(issues during roll/pitch, ascent/descent, sampling in cloud, icing etc.)

10.  Miscellaneous

1.  Hazmat for preflight/postflight calibrations:  Please describe fully any additional hazardous materials - compressed gases, solvents, radioactive ion sources – that you anticipate temporarily bringing onto the aircraft for periodic instrument calibration purposes  (e.g., n-butanol in a CN counter, 210Po in a DMA, a UPS for power, compressed gas cylinders for calibrations, etc.)

none

2.  Fabrication and sheet metal support:  Please describe fully any anticipated requests for fabrication or sheet-metal support during installation in Tampa.  This list should be kept to an absolute minimum; please recognize that this superb AOC resource is quite limited.  To ease the strain on the AOC shop, we will work with each PI to ensure they arrive in Tampa with as much in hand as possible.

none

3.  Ferry flight/check flight procedures.  On occasion, AOC will perform an aircraft check flight, during which the instruments may be flown without power.  Aircraft maintenance needs may also dictate a ferry flight without science crew or SED techs on board.  Instruments should be designed with these eventualities in mind.  However, if your instrument requires standby power during this kind of flight, this may be provided at the discretion of AOC personnel.

If so, the flight crew will need to be briefed well ahead of time to ensure proper instrument operation.  Please provide with this document a bare-minimum checklist of instrument startup and shutdown procedures requested for these flights.  N/A

III.  Ground laboratory space

1.  Tampa space requests: see C3X worksheet

Power requirements: laptop

Special requests:

2.  Field space requests:

Work space, ft2: 10

Number of tables/chairs: 1

Power requirements: laptop

Storage space, ft2: 100 (~3 large Boxes, several small boxes)

Other requests: