NOAA Nitrogen Oxides and Ozone (NOyO3)

Principal investigator:  Tom Ryerson

NOAA Chemical Sciences Division

thomas.b.ryerson@noaa.gov

Co-investigator: Ilana Pollack

University of Colorado-CIRES

ilana.pollack@noaa.gov

The NOAA NOyO3 4-channel chemiluminescence (CL) instrument will provide in-situ measurements of nitric oxide (NO), nitrogen dioxide (NO2), total reactive nitrogen oxides (NOy), and ozone (O3) on the NOAA P-3 during the Southeast Nexus (SENEX) project.  This instrument has flown on the NOAA P-3, the NCAR Electra, and the NASA DC-8 research aircraft on multiple field projects since 1995.  It provides fast-response, chemically specific, high precision, and calibrated measurements of nitrogen oxides and ozone at a spatial resolution of better than 50m at typical P-3 research flight speeds.

Detection is based on the gas-phase CL reaction of NO with O3 at low pressure, resulting in photoemission from electronically excited NO2.  Photons are detected and quantified using pulse counting techniques, providing ~5 to 10 part-per-trillion by volume (pptv) precision at 1 Hz data rates.

One CL channel of the integrated 4-channel instrument is used to measure ambient NO directly, a second channel is equipped with a high-power UV-LED converter to photodissociate ambient NO2 to NO, and a third channel is equipped with a heated gold catalyst to reduce ambient NOy species to NO.  Reagent ozone is added to these sample streams to drive the CL reactions with NO.  Ambient O3 is detected in the fourth channel by adding reagent NO.

Instrument performance is routinely evaluated in flight by standard addition calibrations delivered within a few centimeters of the inlet tips.  The separate NO and NO2 sample paths, detectors, and inlet residence times are identical, permitting artifact-free calculation of ambient NO2 by difference at high time resolution, with no lagging or smoothing relative to NO or to other fast-response measurements aboard the aircraft.  A high-power UV-LED converter developed in our laboratory provides NO2 conversion fractions exceeding 0.6 at a converter sample residence time of 0.11 seconds.  This offers a significant advantage in terms of NO and NO2 spatial resolution compared to other airborne NO2 instruments.  The NOy channel is calibrated to NO, NO2, and HNO3 in flight and the O3 channel is calibrated over an atmospherically-relevant range of ozone mixing ratios in flight.

Measurements of nitrogen oxides and ozone will be used to address several objectives of SENEX including:

- Understanding tropospheric ozone chemistry in the presence of biogenic and anthropogenic precursor emissions

- Understanding air quality – climate interactions

- Comparison with regional air quality models; evaluation of satellite observations

References

Pollack, I.B., Lerner, B.T., and Ryerson, T.B. (2011), Evaluation of ultraviolet light-emitting diodes for detection of atmospheric NO2 by photolysis-chemiluminescence.  Journal of Atmospheric Chemistry 65:111–125, 10.1007/s10874-011-9184-3.

Ryerson, T.B., Williams, E.J., and Fehsenfeld, F.C. (2000), An efficient photolysis system for fast-response NO2 measurements.  Journal of Geophysical Research, 105(D21), 26,447-426,461.

Ryerson, T.B., Huey, L.G., Knapp, K., Neuman, J.A., Parrish, D.D., Sueper, D.T., and Fehsenfeld, F.C. (1999), Design and initial characterization of an inlet for gas-phase NOy measurements from aircraft. Journal of Geophysical Research, 104(D5), 5483-5492.