Month:	________Year:200__													Station: Barrow, Alaska

| Date | CNC butanol level: fill y/n | CN drier flow* ~8.5 lpm | CNC vacuum >14 ” Hg | CNC status lights G/R/Both | PSAP filter change | PSAP flow ~2.0 lpm | Q_stack* ~850 lpm | MFC readout, Neph. ~29 lpm | MFC readout, filters ~29 lpm | Active filter (1-8) | Active filter pressure (60-100 hPa) | Comments (e.g., rain, smoke, instrument issues, vehicle traffic, etc.etc.) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |  |  |  |  |  |
| 21 |  |  |  |  |  |  |  |  |  |  |  |  |
| 22 |  |  |  |  |  |  |  |  |  |  |  |  |
| 23 |  |  |  |  |  |  |  |  |  |  |  |  |
| 24 |  |  |  |  |  |  |  |  |  |  |  |  |
| 25 |  |  |  |  |  |  |  |  |  |  |  |  |
| 26 |  |  |  |  |  |  |  |  |  |  |  |  |
| 27 |  |  |  |  |  |  |  |  |  |  |  |  |
| 28 |  |  |  |  |  |  |  |  |  |  |  |  |
| 29 |  |  |  |  |  |  |  |  |  |  |  |  |
| 30 |  |  |  |  |  |  |  |  |  |  |  |  |
| 31 |  |  |  |  |  |  |  |  |  |  |  |  |

Problems: Contact Anne Jefferson (303) 497-6493 or Betsy Andrews (303) 497-5171			Logging Status: ABCFLMNU

*Check values on computer screen in ‘aerosol window’ <alt+A>

Month:________________  Year: 200___

WEEKLY TASKS

| PSAP ref. filter Date Time (UTC) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Neph span check Date Time (UTC) |  |  |  |  |
| Impactor films Date |  |  |  |  |
| Carousel filters (Filter Rack) Date |  |  |  |  |
| Leak check Date |  |  |  |  |
| Mail carousel filters to PMEL |  |  |  |  |

MONTHLY TASKS

| Mail PSAP filters, checklists to NOAA |  |
| --- | --- |
| Check supply levels | Butanol PSAP filters Impactor films Plastic bags for PSAP filters Neph Lamps Other:___________________ |

## Typical Parameter Values at BRW

| Window | <Alt> key | Parameter | Typical | Lower limit | Upper limit | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Aerosol Sampling Status | A | RH_sample | 40 | 0 | 45 | Check inlet heater |
|  | A | T_sample | 25 | 10 | 40 |  |
|  | A | T_rack | 30 | 20 | 35 | Check air conditioner |
|  | A | T_pumpbox | 32 | 25 | 40 |  |
|  | A | Q_analyzer | 30 | 28 | 32 | Check pump, adjust flow |
| Aerosol dP sensors | A | dP_Neph_imp | 2 – 10 | 2 | 13 | Check for leaks around impactors |
|  | A | dP_CNC_vac | 600 | 550 | 700 | Check CN pump |
|  | A | dP_pitot | 70 | 60 | 80 | Check carbon vane pump |
| CN counter | A | CN flow | 1.4 | 1.3 | 1.5 | Check CN pump |
|  | A | Drier flow | 8.5 | 8 | 9 | Adjust valve, Change drier |
| Best UPS | B | Inverter | OFF |  |  |  |
|  |  | Charger | ON |  |  |  |
|  |  | AC Volts IN | 113 | 110 | 120 | Fix generators |
|  |  | AC Volts OUT | 121 | 118 | 125 |  |
|  |  | AC Amps OUT |  |  |  |  |
|  |  | Output Load VA |  |  |  | Check that all systems are drawing power, and than nothing unauthorized is connected to UPS |
|  |  | Battery DC V | 13.9 | 13 | 14 | Possible UPS malfunction |
|  |  | Frequency | 60 | 50 | 70 | Fix generators |
| PSAP | L | Flow | 2.0 | 1.75 | 2.5 | Adjust flow |
|  | L | Transmittance | 0.9 | 0.7 | 1.0 | Change filter |
|  | L | I_sample | 250000 | 90000 | 300000 | Check light source |
|  | L | I_reference | 570000 | 250000 | 300000 | Check light source |
|  | L | Filter change | FALSE | FALSE | FALSE | You forgot to press <alt><L><M><3> after last filter change. |

| Window | <Alt> | Parameter | Typical | Lower | Upper | Action |
| --- | --- | --- | --- | --- | --- | --- |
| Neph_N | N | Bkg_tot_Blue | 8.0E-6 | 1E-7 | 5*typical |  |
|  |  | Bkg_tot_Green | 8.0E-6 | 1E-7 | 5*typical |  |
|  |  | Bkg_tot_Red | 8.0E-6 | 1E-7 | 5*typical |  |
|  |  | Bkg_back_Blu | 8.0E-7 | 1E-7 | 5*typical |  |
|  |  | Bkg_back_Grn | 8.0E-7 | 1E-7 | 5*typical |  |
|  |  | Bkg_bak_Red | 8.0E-7 | 1E-7 | 5*typical |  |
|  |  | MODE | NBXX | BBXX | ZBXX | Any other value is an error |
|  |  | Neph_press | 1000 | 900 | 1050 |  |
|  |  | SecToGo | 30000 |  |  | Must decrease by one every second |
|  |  | Sample_T | 300 | 290 | 320 | Check room temperature |
|  |  | Inlet_T | 297 | 290 | 320 | Check room temperature |
|  |  | Next_Z | Xx:57:00 |  |  | Possible error in CP.INI entry for [Neph] ZeroAt= entry. |
|  |  | GreenRef | 4.0E+5 | 3.0E+5 | 5.0E+5 | Lamp may be failing. |
|  |  | Lamp_V | 13 | 10 | 14 | Lamp may be failing |
|  |  | Lamp_I | 5.7 | 5 | 7.5 | Lamp_V times Lamp_I should be 75.0 |
| uMAC | U | Room Temp | 38 | 30 | 40 | Check room temperature |
|  |  | Power supply | 5.0 | 4.9 | 5.1 | Power supply bad? |

General commands

| Keystrokes | Action |
| --- | --- |
| <enter><enter> | Brings up message window |
| <control> <M> | Shows root menu |
| <esc> | Shows root menu and hides root menu |
| <alt> <X> | Selects root menu entry ‘X’ when root menu is not focused (X can be A,B,C,F,G,H,L,M,N,O,S,U,V,W) |
| <control><D> | Disc operations menu |
| <control><M><enter> | Online log entry (message window) |
| <F12> | Online log entry (message window) |
| <control><E> | Toggles error window |
| <control><C> | Quit cpdclient |
| <M> | Shows menu for active window |
| <control> <W> or <W> | Closes active window |
| <tab> | Cycles through open windows |
| Arrow keys (in menu) | Move thru menu list |
| Number/letter (in menu) | Quick select on menu list |
| Keystrokes | Action |
| <enter> (in menu/dialog box) | Chooses or toggles selected item |
| <enter> (dialog box) | Begins enter input mode, text will be bolded; pressing <enter> again will end input mode |
| <control><W> (in menu) | Cancels selection |
| <esc> or <control><W> | Exits menu/dialog box without applying changes |
| <enter> (with bottom entry selected in dialog box) | Closes dialog box and applies all changes |

Maintenance commands – note these are ONLY the computer commands needed for maintenance.  The manual describes the maintenance procedures in more detail and should be referred to for exact procedures.

| Keystrokes | Action |
| --- | --- |
| <enter><enter> | Brings up message window |
| <Alt><L><M><3> <Alt><L><M><3> | PSAP filter change Lets computer know psap filter is being changed Lets computer know psap change is over |
| <enter><enter> <Alt><A ><M><3> <Alt> <A><M><3> <Alt> <N><M><5> <enter><enter> | Impactor change Enter message ‘starting impactor film change’ Bypass analyzers Actually change impactors Unbypass analyzers Start Neph zero Enter message ‘end of impactor servicing and neph zero’ |
| <enter><enter> <enter><enter> <enter><enter> | Leak Check Enter message ‘starting leak check’ Enter message ‘ambient values are CN= dp Neph imp =’ Enter message ‘filtered values are CN= dp Neph imp =’ |
| <Alt><F><M,<6> <Alt><F><M><6> | Changing Carousel Filters Selects filter change and computer knows it’s starting Do all the filter change stuff computer knows filter change is over |
| <Alt> <O><M><3> <Alt> <N> | Span check Lets computer know span check is starting, follow prompts Check span check results |