## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
283.2219 | 291.6755 | 281.9926 | 276.7695 | 279.5209
3.701 | 12.1545 | 2.4716 | -2.7514
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
281.2022 | 283.1458 | 277.5162 | 279.7838 | 276.0596
5.1426 | 7.0861 | 1.4566 | 3.7242

## IdD083
Dobson 083 at IZO 2016.Sep.15 Langley Plot Results IUP Constants
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7 | [11 empty] | Correction to
D083 LP adj. A:-1.71 C:-0.16 D:-0.86 Nad:-0.86 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -0.46352 |  | -0.84555 |  | -2.1953 | -1.3497 | 0.38203 | [14 empty] | Limited Mu range
[11 empty] | 17 |  |  |  | 283.2219 | 291.6755 | 281.9926 | 276.7695 | 279.5209 | all Values | -0.16146 |  | -0.85511 |  | -1.7149 | -0.85981 | 0.69366 | [14 empty] | 0.23743 |  | -1.3468
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
137.7 | 77.6 | 246.1 | 15 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 128.9286 | 74.1273 | 233.1302 | 282.3519 | 288.1216 | 280.5733 | 278.1248 | 278.589 | 0.24666 | 0.23155 | 0.25014 | 0.10304 | 0.25363 | 0.50109 | 2016-09-15 07:56 | 4.0541 | 4.2319 | 2016-09-15 07:57 | 3.9977 | 4.1674 | 2016-09-15 07:58 | 3.9428 | 4.1048 | 2016-09-15 07:56 | 4.0257 | 4.1994 | 2016-09-15 07:57 | 3.9701 | 4.1358 | 0.2484 | 0.12852 | 0.25189 | 0.39805
130.6 | 73.2 | 234.5 | 15 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 122.2461 | 70.0991 | 221.1961 | 283.5657 | 288.4017 | 281.1459 | 280.0204 | 279.2473 | 0.2606 | 0.23255 | 0.26409 | 0.10314 | 0.26759 | 0.50211 | 2016-09-15 08:00 | 3.8372 | 3.9853 | 2016-09-15 08:01 | 3.7865 | 3.9282 | 2016-09-15 08:02 | 3.7371 | 3.8727 | 2016-09-15 08:01 | 3.8117 | 3.9565 | 2016-09-15 08:02 | 3.7617 | 3.9003 | 0.26235 | 0.12941 | 0.26584 | 0.39898
123.6 | 69.3 | 223.6 | 15 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 115.7162 | 66.5399 | 210.1047 | 282.9711 | 289.0501 | 281.1655 | 278.4327 | 279.1196 | 0.27457 | 0.23206 | 0.27807 | 0.10337 | 0.28156 | 0.50215 | 2016-09-15 08:04 | 3.642 | 3.7665 | 2016-09-15 08:05 | 3.5963 | 3.7156 | 2016-09-15 08:06 | 3.5516 | 3.6661 | 2016-09-15 08:04 | 3.619 | 3.7409 | 2016-09-15 08:06 | 3.5738 | 3.6907 | 0.27632 | 0.12869 | 0.27981 | 0.39878
117.5 | 65.8 | 213.8 | 15 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 110.0458 | 63.347 | 200.3108 | 283.1196 | 289.7752 | 281.5678 | 278.1095 | 279.449 | 0.28855 | 0.23218 | 0.29204 | 0.10363 | 0.29554 | 0.50287 | 2016-09-15 08:08 | 3.4656 | 3.5711 | 2016-09-15 08:09 | 3.4241 | 3.5255 | 2016-09-15 08:10 | 3.3837 | 3.4811 | 2016-09-15 08:08 | 3.4447 | 3.5481 | 2016-09-15 08:10 | 3.4038 | 3.5031 | 0.2903 | 0.12855 | 0.29379 | 0.39924
111.9 | 62.7 | 204.7 | 15 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.8596 | 60.5032 | 191.2887 | 283.0569 | 290.8591 | 281.7349 | 277.1394 | 279.3933 | 0.30252 | 0.23213 | 0.30601 | 0.10402 | 0.3095 | 0.50317 | 2016-09-15 08:12 | 3.3056 | 3.3956 | 2016-09-15 08:13 | 3.2679 | 3.3545 | 2016-09-15 08:14 | 3.231 | 3.3145 | 2016-09-15 08:12 | 3.2866 | 3.375 | 2016-09-15 08:13 | 3.2493 | 3.3344 | 0.30427 | 0.12812 | 0.30775 | 0.39915
106.9 | 59.7 | 196.1 | 15 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 100.2264 | 57.7387 | 182.8238 | 283.2909 | 290.2981 | 281.4529 | 277.9708 | 279.1865 | 0.31647 | 0.23232 | 0.31995 | 0.10382 | 0.32343 | 0.50266 | 2016-09-15 08:16 | 3.1599 | 3.2373 | 2016-09-15 08:17 | 3.1255 | 3.2001 | 2016-09-15 08:18 | 3.0918 | 3.1638 | 2016-09-15 08:17 | 3.1426 | 3.2186 | 2016-09-15 08:18 | 3.1086 | 3.1818 | 0.31821 | 0.12851 | 0.32169 | 0.39885
102.2 | 57.3 | 188.8 | 15 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 95.8623 | 55.5154 | 175.6299 | 282.9562 | 292.5007 | 282.1961 | 275.66 | 279.5688 | 0.33038 | 0.23205 | 0.33386 | 0.1046 | 0.33733 | 0.50399 | 2016-09-15 08:20 | 3.0268 | 3.0937 | 2016-09-15 08:21 | 2.9953 | 3.0599 | 2016-09-15 08:22 | 2.9645 | 3.0269 | 2016-09-15 08:20 | 3.011 | 3.0767 | 2016-09-15 08:22 | 2.9798 | 3.0433 | 0.33212 | 0.12745 | 0.33559 | 0.39939
98.1 | 55 | 181.5 | 15 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 92.0499 | 53.376 | 168.4293 | 283.376 | 293.6554 | 281.7466 | 275.4997 | 278.7215 | 0.34425 | 0.23239 | 0.34771 | 0.10502 | 0.35117 | 0.50319 | 2016-09-15 08:24 | 2.9048 | 2.9631 | 2016-09-15 08:25 | 2.8759 | 2.9322 | 2016-09-15 08:26 | 2.8476 | 2.902 | 2016-09-15 08:24 | 2.8903 | 2.9475 | 2016-09-15 08:26 | 2.8617 | 2.917 | 0.34598 | 0.12738 | 0.34944 | 0.39817
94.2 | 52.7 | 175.1 | 15 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 88.4343 | 51.2218 | 162.1584 | 283.2809 | 292.8669 | 282.0157 | 275.9307 | 279.2585 | 0.35807 | 0.23232 | 0.36152 | 0.10473 | 0.36496 | 0.50367 | 2016-09-15 08:28 | 2.7927 | 2.8436 | 2016-09-15 08:29 | 2.7661 | 2.8154 | 2016-09-15 08:30 | 2.7401 | 2.7877 | 2016-09-15 08:28 | 2.7794 | 2.8294 | 2016-09-15 08:29 | 2.753 | 2.8015 | 0.35979 | 0.12758 | 0.36324 | 0.39893
90.6 | 50.6 | 169 | 15 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 85.0985 | 49.2395 | 156.1971 | 283.1488 | 292.0843 | 281.9045 | 276.2932 | 279.3181 | 0.37182 | 0.23221 | 0.37525 | 0.10445 | 0.37867 | 0.50347 | 2016-09-15 08:32 | 2.6895 | 2.7341 | 2016-09-15 08:33 | 2.6649 | 2.7082 | 2016-09-15 08:34 | 2.6408 | 2.6827 | 2016-09-15 08:33 | 2.6771 | 2.7211 | 2016-09-15 08:34 | 2.6528 | 2.6954 | 0.37354 | 0.12775 | 0.37696 | 0.39901
87.6 | 48.8 | 163.5 | 15 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 82.3155 | 47.53 | 150.8145 | 284.407 | 292.6096 | 282.0768 | 278.1107 | 279.4044 | 0.3855 | 0.23324 | 0.38891 | 0.10464 | 0.39231 | 0.50378 | 2016-09-15 08:36 | 2.594 | 2.6334 | 2016-09-15 08:37 | 2.5713 | 2.6095 | 2016-09-15 08:38 | 2.549 | 2.586 | 2016-09-15 08:36 | 2.5826 | 2.6214 | 2016-09-15 08:38 | 2.5601 | 2.5977 | 0.38721 | 0.1286 | 0.39061 | 0.39913
84.2 | 47.1 | 158.4 | 15 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 79.171 | 45.906 | 145.8492 | 282.8839 | 292.6534 | 282.3004 | 275.3699 | 279.6749 | 0.3991 | 0.23199 | 0.40249 | 0.10466 | 0.40587 | 0.50418 | 2016-09-15 08:40 | 2.5056 | 2.5405 | 2016-09-15 08:41 | 2.4845 | 2.5184 | 2016-09-15 08:42 | 2.4638 | 2.4967 | 2016-09-15 08:40 | 2.495 | 2.5294 | 2016-09-15 08:42 | 2.4741 | 2.5075 | 0.4008 | 0.12733 | 0.40418 | 0.39952
81.4 | 45.5 | 153.6 | 15 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 76.594 | 44.3688 | 141.2031 | 283.0793 | 292.3342 | 282.4426 | 275.9579 | 279.9343 | 0.41262 | 0.23215 | 0.41598 | 0.10454 | 0.41934 | 0.50443 | 2016-09-15 08:44 | 2.4235 | 2.4545 | 2016-09-15 08:45 | 2.4039 | 2.434 | 2016-09-15 08:46 | 2.3847 | 2.4139 | 2016-09-15 08:44 | 2.4137 | 2.4442 | 2016-09-15 08:46 | 2.3943 | 2.4239 | 0.4143 | 0.12761 | 0.41766 | 0.39989
78.8 | 44.2 | 149.2 | 15 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 74.2026 | 43.1138 | 136.9378 | 283.2933 | 294.1063 | 282.7272 | 274.9624 | 279.8479 | 0.42604 | 0.23233 | 0.42938 | 0.10518 | 0.43271 | 0.50494 | 2016-09-15 08:48 | 2.3472 | 2.3748 | 2016-09-15 08:49 | 2.3289 | 2.3558 | 2016-09-15 08:50 | 2.311 | 2.3371 | 2016-09-15 08:49 | 2.338 | 2.3653 | 2016-09-15 08:50 | 2.3199 | 2.3464 | 0.42771 | 0.12715 | 0.43105 | 0.39976
76.3 | 42.6 | 144.9 | 15 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 71.8995 | 41.5617 | 132.8131 | 283.1005 | 291.1195 | 282.6433 | 276.926 | 280.4922 | 0.43936 | 0.23217 | 0.44268 | 0.10411 | 0.44598 | 0.50479 | 2016-09-15 08:52 | 2.276 | 2.3007 | 2016-09-15 08:53 | 2.259 | 2.283 | 2016-09-15 08:54 | 2.2422 | 2.2656 | 2016-09-15 08:52 | 2.2675 | 2.2918 | 2016-09-15 08:54 | 2.2506 | 2.2743 | 0.44102 | 0.12806 | 0.44433 | 0.40068
74.1 | 41.5 | 141.1 | 15 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 69.8721 | 40.4895 | 129.1852 | 283.5816 | 292.9569 | 283.1414 | 276.3547 | 280.656 | 0.45258 | 0.23256 | 0.45586 | 0.10477 | 0.45914 | 0.50568 | 2016-09-15 08:56 | 2.2096 | 2.2317 | 2016-09-15 08:57 | 2.1936 | 2.2152 | 2016-09-15 08:58 | 2.178 | 2.199 | 2016-09-15 08:56 | 2.2016 | 2.2234 | 2016-09-15 08:58 | 2.1858 | 2.2071 | 0.45422 | 0.1278 | 0.4575 | 0.40091
71.9 | 40.5 | 137.3 | 15 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 67.8447 | 39.5109 | 125.5642 | 283.3095 | 295.0903 | 283.0389 | 274.2194 | 279.9947 | 0.46568 | 0.23234 | 0.46894 | 0.10553 | 0.47219 | 0.50549 | 2016-09-15 09:00 | 2.1474 | 2.1673 | 2016-09-15 09:01 | 2.1325 | 2.1519 | 2016-09-15 09:02 | 2.1178 | 2.1367 | 2016-09-15 09:00 | 2.1399 | 2.1596 | 2016-09-15 09:02 | 2.1251 | 2.1443 | 0.46731 | 0.12681 | 0.47056 | 0.39996

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.15 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:0.48 C:0.30 D:-0.01 Nad:0.49 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 281.2022 | 283.1458 | 277.5162 | 279.7838 | 276.0596 | All Obs | 0.30206 |  | -0.0095629 |  | 0.48036 | 0.48993 | 0.31162
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
137.7 | 77.6 | 246.1 | 15 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 128.4651 | 73.2817 | 230.9349 | 280.9578 | 282.2073 | 277.4557 | 280.226 | 276.1718 | 0.24666 | 0.23041 | 0.25014 | 0.10092 | 0.25363 | 0.49552 | 2016-09-15 07:56 | 4.0541 | 4.2319 | 2016-09-15 07:57 | 3.9977 | 4.1674 | 2016-09-15 07:58 | 3.9428 | 4.1048 | 2016-09-15 07:56 | 4.0257 | 4.1994 | 2016-09-15 07:57 | 3.9701 | 4.1358
130.6 | 73.2 | 234.5 | 15 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 121.7826 | 69.2536 | 219.0008 | 282.0928 | 282.1575 | 277.8567 | 282.2369 | 276.6982 | 0.2606 | 0.23134 | 0.26409 | 0.1009 | 0.26759 | 0.49624 | 2016-09-15 08:00 | 3.8372 | 3.9853 | 2016-09-15 08:01 | 3.7865 | 3.9282 | 2016-09-15 08:02 | 3.7371 | 3.8727 | 2016-09-15 08:01 | 3.8117 | 3.9565 | 2016-09-15 08:02 | 3.7617 | 3.9003
123.6 | 69.3 | 223.6 | 15 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 115.2527 | 65.6943 | 207.9094 | 281.4192 | 282.4755 | 277.7046 | 280.7646 | 276.4384 | 0.27457 | 0.23079 | 0.27807 | 0.10102 | 0.28156 | 0.49597 | 2016-09-15 08:04 | 3.642 | 3.7665 | 2016-09-15 08:05 | 3.5963 | 3.7156 | 2016-09-15 08:06 | 3.5516 | 3.6661 | 2016-09-15 08:04 | 3.619 | 3.7409 | 2016-09-15 08:06 | 3.5738 | 3.6907
117.5 | 65.8 | 213.8 | 15 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 109.5823 | 62.5014 | 198.1155 | 281.4887 | 282.8701 | 277.935 | 280.5569 | 276.6358 | 0.28855 | 0.23085 | 0.29204 | 0.10116 | 0.29554 | 0.49638 | 2016-09-15 08:08 | 3.4656 | 3.5711 | 2016-09-15 08:09 | 3.4241 | 3.5255 | 2016-09-15 08:10 | 3.3837 | 3.4811 | 2016-09-15 08:08 | 3.4447 | 3.5481 | 2016-09-15 08:10 | 3.4038 | 3.5031
111.9 | 62.7 | 204.7 | 15 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.396 | 59.6576 | 189.0934 | 281.3471 | 283.6238 | 277.9306 | 279.7021 | 276.4481 | 0.30252 | 0.23073 | 0.30601 | 0.10143 | 0.3095 | 0.49637 | 2016-09-15 08:12 | 3.3056 | 3.3956 | 2016-09-15 08:13 | 3.2679 | 3.3545 | 2016-09-15 08:14 | 3.231 | 3.3145 | 2016-09-15 08:12 | 3.2866 | 3.375 | 2016-09-15 08:13 | 3.2493 | 3.3344
106.9 | 59.7 | 196.1 | 15 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 99.7629 | 56.8932 | 180.6285 | 281.5022 | 282.7333 | 277.4773 | 280.6486 | 276.1097 | 0.31647 | 0.23086 | 0.31995 | 0.10111 | 0.32343 | 0.49556 | 2016-09-15 08:16 | 3.1599 | 3.2373 | 2016-09-15 08:17 | 3.1255 | 3.2001 | 2016-09-15 08:18 | 3.0918 | 3.1638 | 2016-09-15 08:17 | 3.1426 | 3.2186 | 2016-09-15 08:18 | 3.1086 | 3.1818
102.2 | 57.3 | 188.8 | 15 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 95.3988 | 54.6698 | 173.4347 | 281.0889 | 284.6071 | 278.0498 | 278.4525 | 276.3609 | 0.33038 | 0.23052 | 0.33386 | 0.10178 | 0.33733 | 0.49658 | 2016-09-15 08:20 | 3.0268 | 3.0937 | 2016-09-15 08:21 | 2.9953 | 3.0599 | 2016-09-15 08:22 | 2.9645 | 3.0269 | 2016-09-15 08:20 | 3.011 | 3.0767 | 2016-09-15 08:22 | 2.9798 | 3.0433
98.1 | 55 | 181.5 | 15 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 91.5864 | 52.5304 | 166.234 | 281.4303 | 285.4341 | 277.4301 | 278.4066 | 275.3828 | 0.34425 | 0.2308 | 0.34771 | 0.10208 | 0.35117 | 0.49548 | 2016-09-15 08:24 | 2.9048 | 2.9631 | 2016-09-15 08:25 | 2.8759 | 2.9322 | 2016-09-15 08:26 | 2.8476 | 2.902 | 2016-09-15 08:24 | 2.8903 | 2.9475 | 2016-09-15 08:26 | 2.8617 | 2.917
94.2 | 52.7 | 175.1 | 15 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 87.9708 | 50.3762 | 159.9631 | 281.2571 | 284.3193 | 277.5297 | 278.9515 | 275.7896 | 0.35807 | 0.23066 | 0.36152 | 0.10168 | 0.36496 | 0.49566 | 2016-09-15 08:28 | 2.7927 | 2.8436 | 2016-09-15 08:29 | 2.7661 | 2.8154 | 2016-09-15 08:30 | 2.7401 | 2.7877 | 2016-09-15 08:28 | 2.7794 | 2.8294 | 2016-09-15 08:29 | 2.753 | 2.8015
90.6 | 50.6 | 169 | 15 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 84.6349 | 48.3939 | 154.0019 | 281.0473 | 283.2119 | 277.2499 | 279.4272 | 275.7198 | 0.37182 | 0.23048 | 0.37525 | 0.10128 | 0.37867 | 0.49516 | 2016-09-15 08:32 | 2.6895 | 2.7341 | 2016-09-15 08:33 | 2.6649 | 2.7082 | 2016-09-15 08:34 | 2.6408 | 2.6827 | 2016-09-15 08:33 | 2.6771 | 2.7211 | 2016-09-15 08:34 | 2.6528 | 2.6954
87.6 | 48.8 | 163.5 | 15 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 81.852 | 46.6845 | 148.6193 | 282.2281 | 283.4143 | 277.2545 | 281.3574 | 275.6772 | 0.3855 | 0.23145 | 0.38891 | 0.10135 | 0.39231 | 0.49516 | 2016-09-15 08:36 | 2.594 | 2.6334 | 2016-09-15 08:37 | 2.5713 | 2.6095 | 2016-09-15 08:38 | 2.549 | 2.586 | 2016-09-15 08:36 | 2.5826 | 2.6214 | 2016-09-15 08:38 | 2.5601 | 2.5977
84.2 | 47.1 | 158.4 | 15 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 78.7075 | 45.0605 | 143.6539 | 280.6282 | 283.1369 | 277.3115 | 278.7286 | 275.8198 | 0.3991 | 0.23014 | 0.40249 | 0.10126 | 0.40587 | 0.49527 | 2016-09-15 08:40 | 2.5056 | 2.5405 | 2016-09-15 08:41 | 2.4845 | 2.5184 | 2016-09-15 08:42 | 2.4638 | 2.4967 | 2016-09-15 08:40 | 2.495 | 2.5294 | 2016-09-15 08:42 | 2.4741 | 2.5075
81.4 | 45.5 | 153.6 | 15 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 76.1305 | 43.5232 | 139.0078 | 280.7472 | 282.4987 | 277.2881 | 279.4278 | 275.952 | 0.41262 | 0.23024 | 0.41598 | 0.10103 | 0.41934 | 0.49522 | 2016-09-15 08:44 | 2.4235 | 2.4545 | 2016-09-15 08:45 | 2.4039 | 2.434 | 2016-09-15 08:46 | 2.3847 | 2.4139 | 2016-09-15 08:44 | 2.4137 | 2.4442 | 2016-09-15 08:46 | 2.3943 | 2.4239
78.8 | 44.2 | 149.2 | 15 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 73.7391 | 42.2683 | 134.7425 | 280.8853 | 283.9541 | 277.4083 | 278.5427 | 275.7395 | 0.42604 | 0.23035 | 0.42938 | 0.10155 | 0.43271 | 0.49544 | 2016-09-15 08:48 | 2.3472 | 2.3748 | 2016-09-15 08:49 | 2.3289 | 2.3558 | 2016-09-15 08:50 | 2.311 | 2.3371 | 2016-09-15 08:49 | 2.338 | 2.3653 | 2016-09-15 08:50 | 2.3199 | 2.3464
76.3 | 42.6 | 144.9 | 15 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 71.436 | 40.7161 | 130.6178 | 280.6172 | 280.6529 | 277.1613 | 280.6159 | 276.2584 | 0.43936 | 0.23013 | 0.44268 | 0.10037 | 0.44598 | 0.495 | 2016-09-15 08:52 | 2.276 | 2.3007 | 2016-09-15 08:53 | 2.259 | 2.283 | 2016-09-15 08:54 | 2.2422 | 2.2656 | 2016-09-15 08:52 | 2.2675 | 2.2918 | 2016-09-15 08:54 | 2.2506 | 2.2743
74.1 | 41.5 | 141.1 | 15 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 69.4086 | 39.6439 | 126.9899 | 281.0236 | 282.1785 | 277.4976 | 280.1533 | 276.2981 | 0.45258 | 0.23046 | 0.45586 | 0.10091 | 0.45914 | 0.4956 | 2016-09-15 08:56 | 2.2096 | 2.2317 | 2016-09-15 08:57 | 2.1936 | 2.2152 | 2016-09-15 08:58 | 2.178 | 2.199 | 2016-09-15 08:56 | 2.2016 | 2.2234 | 2016-09-15 08:58 | 2.1858 | 2.2071
71.9 | 40.5 | 137.3 | 15 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 67.3812 | 38.6653 | 123.3689 | 280.6774 | 284.0028 | 277.2348 | 278.1257 | 275.5137 | 0.46568 | 0.23018 | 0.46894 | 0.10156 | 0.47219 | 0.49513 | 2016-09-15 09:00 | 2.1474 | 2.1673 | 2016-09-15 09:01 | 2.1325 | 2.1519 | 2016-09-15 09:02 | 2.1178 | 2.1367 | 2016-09-15 09:00 | 2.1399 | 2.1596 | 2016-09-15 09:02 | 2.1251 | 2.1443

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.786 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.82009 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.35762 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.4284 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46249 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## Comp Page
Komyhr adjusted |  |  |  |  | IUPmultiply |  |  |  |  | B&P B multiply |  |  |  |  | D&M multiply
278.832 | 278.9001 | 278.8625 | 278.873 | 278.8221 | 276.6571 | 268.9155 | 276.0181 | 283.0561 | 277.8463 | 278.358 | 274.0194 | 277.998 | 281.9964 | 278.9894 | 274.5101 | 266.6843 | 275.7671 | 280.9924 | 278.1251
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
277.9755 | 275.5019 | 277.4589 | 280.2386 | 277.8925 | 280.1607 | 285.7311 | 280.3182 | 276.0971 | 278.8686 | 278.4488 | 280.4091 | 278.3217 | 277.1347 | 277.7259 | 282.3519 | 288.1216 | 280.5733 | 278.1248 | 278.589
279.1705 | 275.7697 | 278.0252 | 282.1485 | 278.5492 | 281.3651 | 286.0088 | 280.8902 | 277.9789 | 279.5276 | 279.6459 | 280.6816 | 278.8897 | 279.0235 | 278.3822 | 283.5657 | 288.4017 | 281.1459 | 280.0204 | 279.2473
278.5851 | 276.3897 | 278.0446 | 280.5488 | 278.4218 | 280.7751 | 286.6519 | 280.9099 | 276.4027 | 279.3997 | 279.0595 | 281.3127 | 278.9092 | 277.4414 | 278.2548 | 282.9711 | 289.0501 | 281.1655 | 278.4327 | 279.1196
278.7312 | 277.083 | 278.4424 | 280.2231 | 278.7504 | 280.9224 | 287.3709 | 281.3117 | 276.0819 | 279.7295 | 279.2059 | 282.0183 | 279.3082 | 277.1194 | 278.5833 | 283.1196 | 289.7752 | 281.5678 | 278.1095 | 279.449
278.6695 | 278.1194 | 278.6077 | 279.2456 | 278.6948 | 280.8602 | 288.4458 | 281.4787 | 275.1188 | 279.6736 | 279.1441 | 283.0732 | 279.474 | 276.1527 | 278.5277 | 283.0569 | 290.8591 | 281.7349 | 277.1394 | 279.3933
278.8999 | 277.5831 | 278.3287 | 280.0834 | 278.4885 | 281.0924 | 287.8895 | 281.1969 | 275.9442 | 279.4666 | 279.3748 | 282.5273 | 279.1942 | 276.9812 | 278.3215 | 283.2909 | 290.2981 | 281.4529 | 277.9708 | 279.1865
278.5704 | 279.6892 | 279.0638 | 277.755 | 278.8699 | 280.7603 | 290.0739 | 281.9395 | 273.6502 | 279.8494 | 279.0448 | 284.671 | 279.9316 | 274.6786 | 278.7027 | 282.9562 | 292.5007 | 282.1961 | 275.66 | 279.5688
278.9837 | 280.7933 | 278.6193 | 277.5935 | 278.0247 | 281.1769 | 291.2189 | 281.4905 | 273.4911 | 279.0012 | 279.4588 | 285.7947 | 279.4857 | 274.5189 | 277.858 | 283.376 | 293.6554 | 281.7466 | 275.4997 | 278.7215
278.89 | 280.0393 | 278.8853 | 278.0278 | 278.5603 | 281.0825 | 290.437 | 281.7592 | 273.919 | 279.5387 | 279.365 | 285.0273 | 279.7525 | 274.9484 | 278.3933 | 283.2809 | 292.8669 | 282.0157 | 275.9307 | 279.2585
278.76 | 279.291 | 278.7754 | 278.393 | 278.6199 | 280.9514 | 289.6609 | 281.6482 | 274.2788 | 279.5984 | 279.2347 | 284.2656 | 279.6423 | 275.3095 | 278.4528 | 283.1488 | 292.0843 | 281.9045 | 276.2932 | 279.3181
279.9987 | 279.7933 | 278.9458 | 280.2243 | 278.7059 | 282.1998 | 290.1818 | 281.8204 | 276.0831 | 279.6848 | 280.4755 | 284.7769 | 279.8132 | 277.1206 | 278.5387 | 284.407 | 292.6096 | 282.0768 | 278.1107 | 279.4044
278.4992 | 279.8352 | 279.1669 | 277.4627 | 278.9757 | 280.6886 | 290.2253 | 282.0437 | 273.3623 | 279.9556 | 278.9735 | 284.8195 | 280.035 | 274.3896 | 278.8085 | 282.8839 | 292.6534 | 282.3004 | 275.3699 | 279.6749
278.6916 | 279.5299 | 279.3075 | 278.0551 | 279.2345 | 280.8825 | 289.9087 | 282.1858 | 273.946 | 280.2152 | 279.1662 | 284.5088 | 280.176 | 274.9754 | 279.067 | 283.0793 | 292.3342 | 282.4426 | 275.9579 | 279.9343
278.9023 | 281.2244 | 279.5889 | 277.0521 | 279.1483 | 281.0948 | 291.6661 | 282.4701 | 272.9577 | 280.1288 | 279.3772 | 286.2335 | 280.4583 | 273.9835 | 278.9809 | 283.2933 | 294.1063 | 282.7272 | 274.9624 | 279.8479
278.7124 | 278.3685 | 279.506 | 279.0306 | 279.7909 | 280.9035 | 288.7041 | 282.3863 | 274.907 | 280.7736 | 279.1871 | 283.3267 | 280.3751 | 275.9401 | 279.6232 | 283.1005 | 291.1195 | 282.6433 | 276.926 | 280.4922
279.1861 | 280.1254 | 279.9985 | 278.455 | 279.9544 | 281.3808 | 290.5262 | 282.8839 | 274.3399 | 280.9377 | 279.6615 | 285.1149 | 280.8692 | 275.3709 | 279.7865 | 283.5816 | 292.9569 | 283.1414 | 276.3547 | 280.656
278.9182 | 282.1654 | 279.8971 | 276.3035 | 279.2947 | 281.1108 | 292.642 | 282.7815 | 272.2202 | 280.2757 | 279.3931 | 287.1912 | 280.7675 | 273.2432 | 279.1273 | 283.3095 | 295.0903 | 283.0389 | 274.2194 | 279.9947