## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
278.7668 | 292.7235 | 279.0781 | 268.0241 | 275.6078
3.159 | 17.1157 | 3.4703 | -7.5837
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
274.1912 | 276.6816 | 275.515 | 272.3238 | 275.1982
-1.007 | 1.4834 | 0.31679 | -2.8744

## IdD083
Dobson 083 at IZO 2016.Sep.18 Langley Plot Results
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7 | [11 empty] | Correction to
D083 LP adj. A:-1.24 C:-0.90 D:-0.79 Nad:-0.45 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -0.77485 |  | -1.1779 |  | -1.2889 | -0.11101 | 0.40307 | [14 empty] | Limited Mu range
[11 empty] | 34 |  |  |  | 278.7668 | 292.7235 | 279.0781 | 268.0241 | 275.6078 | all Values | -0.90352 |  | -0.78568 |  | -1.2366 | -0.45092 | -0.11784 | [14 empty] | 0.39997 |  | -0.18767
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
138.6 | 79.5 | 248.1 | 18 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 129.7752 | 75.8698 | 235.1604 | 274.5551 | 286.7356 | 276.4415 | 265.3915 | 273.7532 | 0.24162 | 0.22692 | 0.24509 | 0.1034 | 0.24856 | 0.49416 | 2016-09-18 07:56 | 4.1387 | 4.3291 | 2016-09-18 07:57 | 4.0801 | 4.2617 | 2016-09-18 07:58 | 4.0231 | 4.1964 | 2016-09-18 07:56 | 4.1092 | 4.2951 | 2016-09-18 07:57 | 4.0514 | 4.2288 | 0.24335 | 0.12352 | 0.24683 | 0.39076
131.1 | 74.9 | 236.3 | 18 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 122.7152 | 71.6544 | 223.0587 | 275.1135 | 286.8954 | 277.2281 | 266.2149 | 274.7119 | 0.25552 | 0.22738 | 0.259 | 0.10346 | 0.26248 | 0.49557 | 2016-09-18 08:00 | 3.9136 | 4.0717 | 2016-09-18 08:01 | 3.861 | 4.0122 | 2016-09-18 08:02 | 3.8098 | 3.9544 | 2016-09-18 08:01 | 3.8871 | 4.0417 | 2016-09-18 08:02 | 3.8352 | 3.9831 | 0.25726 | 0.12392 | 0.26074 | 0.39211
124.8 | 71 | 225.3 | 18 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 116.8335 | 68.0892 | 211.8125 | 277.0909 | 288.5848 | 277.4392 | 268.381 | 274.5602 | 0.26945 | 0.22902 | 0.27293 | 0.10407 | 0.27642 | 0.49594 | 2016-09-18 08:04 | 3.7113 | 3.8438 | 2016-09-18 08:05 | 3.6639 | 3.7909 | 2016-09-18 08:06 | 3.6177 | 3.7394 | 2016-09-18 08:04 | 3.6874 | 3.8171 | 2016-09-18 08:06 | 3.6406 | 3.765 | 0.27119 | 0.12495 | 0.27468 | 0.39188
118.5 | 67.3 | 215.5 | 18 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 110.974 | 64.7172 | 202.0106 | 277.1049 | 288.9182 | 278.2142 | 268.1202 | 275.4555 | 0.28339 | 0.22903 | 0.28688 | 0.10419 | 0.29036 | 0.49733 | 2016-09-18 08:08 | 3.5287 | 3.6407 | 2016-09-18 08:09 | 3.4858 | 3.5934 | 2016-09-18 08:10 | 3.444 | 3.5473 | 2016-09-18 08:08 | 3.5071 | 3.6169 | 2016-09-18 08:10 | 3.4648 | 3.5702 | 0.28513 | 0.12484 | 0.28862 | 0.39314
112.6 | 64.2 | 206.1 | 18 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 105.5077 | 61.8805 | 192.6674 | 276.4918 | 290.9718 | 278.2022 | 265.4207 | 274.9286 | 0.29733 | 0.22852 | 0.30081 | 0.10493 | 0.30429 | 0.49731 | 2016-09-18 08:12 | 3.3633 | 3.4587 | 2016-09-18 08:13 | 3.3244 | 3.4162 | 2016-09-18 08:14 | 3.2863 | 3.3747 | 2016-09-18 08:12 | 3.3437 | 3.4373 | 2016-09-18 08:13 | 3.3052 | 3.3953 | 0.29907 | 0.12359 | 0.30255 | 0.39238
107.5 | 61.2 | 197.4 | 18 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 100.7826 | 59.1229 | 184.1048 | 276.7551 | 291.3876 | 278.0758 | 265.5477 | 274.6709 | 0.31125 | 0.22874 | 0.31472 | 0.10508 | 0.31819 | 0.49708 | 2016-09-18 08:16 | 3.2129 | 3.2948 | 2016-09-18 08:17 | 3.1774 | 3.2563 | 2016-09-18 08:18 | 3.1427 | 3.2188 | 2016-09-18 08:17 | 3.1951 | 3.2754 | 2016-09-18 08:18 | 3.16 | 3.2374 | 0.31298 | 0.12366 | 0.31646 | 0.392
102.7 | 58.3 | 189.4 | 18 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 96.3274 | 56.4427 | 176.2206 | 276.4115 | 290.2433 | 277.823 | 265.8088 | 274.6483 | 0.32513 | 0.22845 | 0.3286 | 0.10466 | 0.33206 | 0.49663 | 2016-09-18 08:20 | 3.0757 | 3.1464 | 2016-09-18 08:21 | 3.0432 | 3.1114 | 2016-09-18 08:22 | 3.0115 | 3.0773 | 2016-09-18 08:20 | 3.0594 | 3.1288 | 2016-09-18 08:22 | 3.0273 | 3.0943 | 0.32686 | 0.12379 | 0.33033 | 0.39197
98.5 | 55.7 | 182.4 | 18 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 92.4214 | 54.0284 | 169.3181 | 276.7426 | 289.4247 | 278.1861 | 267.0172 | 275.314 | 0.33897 | 0.22873 | 0.34243 | 0.10437 | 0.34587 | 0.49728 | 2016-09-18 08:24 | 2.9501 | 3.0115 | 2016-09-18 08:25 | 2.9203 | 2.9796 | 2016-09-18 08:26 | 2.8912 | 2.9485 | 2016-09-18 08:24 | 2.9351 | 2.9955 | 2016-09-18 08:26 | 2.9057 | 2.964 | 0.3407 | 0.12436 | 0.34415 | 0.39291
94.8 | 53.6 | 175.7 | 18 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 88.9901 | 52.0669 | 162.7437 | 277.7062 | 291.0899 | 278.0769 | 267.4272 | 274.7598 | 0.35276 | 0.22952 | 0.3562 | 0.10497 | 0.35963 | 0.49708 | 2016-09-18 08:28 | 2.8348 | 2.8883 | 2016-09-18 08:29 | 2.8074 | 2.8592 | 2016-09-18 08:30 | 2.7806 | 2.8308 | 2016-09-18 08:28 | 2.821 | 2.8737 | 2016-09-18 08:29 | 2.794 | 2.8449 | 0.35448 | 0.12455 | 0.35792 | 0.39211
91.1 | 51.6 | 170 | 18 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 85.5621 | 50.185 | 157.1749 | 277.4517 | 291.9341 | 278.9866 | 266.3135 | 275.6889 | 0.36649 | 0.22931 | 0.36991 | 0.10527 | 0.37332 | 0.49871 | 2016-09-18 08:32 | 2.7286 | 2.7756 | 2016-09-18 08:33 | 2.7034 | 2.7489 | 2016-09-18 08:34 | 2.6787 | 2.7227 | 2016-09-18 08:33 | 2.7159 | 2.7622 | 2016-09-18 08:34 | 2.691 | 2.7357 | 0.3682 | 0.12404 | 0.37161 | 0.39344
87.9 | 49.7 | 164.1 | 18 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 82.5937 | 48.386 | 151.402 | 278.0816 | 292.0898 | 278.5288 | 267.3027 | 275.0785 | 0.38014 | 0.22984 | 0.38354 | 0.10533 | 0.38694 | 0.49789 | 2016-09-18 08:36 | 2.6306 | 2.672 | 2016-09-18 08:37 | 2.6073 | 2.6474 | 2016-09-18 08:38 | 2.5844 | 2.6233 | 2016-09-18 08:36 | 2.6189 | 2.6596 | 2016-09-18 08:38 | 2.5958 | 2.6353 | 0.38184 | 0.1245 | 0.38524 | 0.39256
85 | 47.8 | 159 | 18 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 79.9092 | 46.5759 | 146.4302 | 278.9898 | 290.623 | 278.9111 | 270.0428 | 275.9302 | 0.39371 | 0.23059 | 0.39709 | 0.1048 | 0.40047 | 0.49858 | 2016-09-18 08:40 | 2.5399 | 2.5765 | 2016-09-18 08:41 | 2.5183 | 2.5538 | 2016-09-18 08:42 | 2.4971 | 2.5316 | 2016-09-18 08:40 | 2.5291 | 2.5651 | 2016-09-18 08:42 | 2.5077 | 2.5426 | 0.3954 | 0.12578 | 0.39878 | 0.39377
82.1 | 46.3 | 154.1 | 18 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 77.2377 | 45.1384 | 141.6881 | 278.9852 | 291.8978 | 278.9947 | 269.0423 | 275.7145 | 0.4072 | 0.23058 | 0.41056 | 0.10526 | 0.41391 | 0.49872 | 2016-09-18 08:44 | 2.4558 | 2.4883 | 2016-09-18 08:45 | 2.4357 | 2.4673 | 2016-09-18 08:46 | 2.416 | 2.4467 | 2016-09-18 08:44 | 2.4457 | 2.4777 | 2016-09-18 08:46 | 2.4258 | 2.4569 | 0.40888 | 0.12532 | 0.41223 | 0.39346
79.4 | 44.7 | 149.4 | 18 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 74.7548 | 43.5971 | 137.1309 | 278.9716 | 290.7371 | 278.727 | 269.9118 | 275.674 | 0.42059 | 0.23057 | 0.42392 | 0.10484 | 0.42725 | 0.49825 | 2016-09-18 08:48 | 2.3776 | 2.4066 | 2016-09-18 08:49 | 2.3589 | 2.387 | 2016-09-18 08:50 | 2.3406 | 2.3679 | 2016-09-18 08:49 | 2.3682 | 2.3968 | 2016-09-18 08:50 | 2.3497 | 2.3774 | 0.42226 | 0.12573 | 0.42559 | 0.3934
76.7 | 43.3 | 145.1 | 18 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 72.2681 | 42.2418 | 133.0042 | 278.0351 | 290.5474 | 278.7582 | 268.393 | 275.7624 | 0.43388 | 0.2298 | 0.43719 | 0.10477 | 0.44049 | 0.4983 | 2016-09-18 08:52 | 2.3048 | 2.3306 | 2016-09-18 08:53 | 2.2873 | 2.3125 | 2016-09-18 08:54 | 2.2702 | 2.2947 | 2016-09-18 08:52 | 2.296 | 2.3215 | 2016-09-18 08:54 | 2.2787 | 2.3035 | 0.43554 | 0.12502 | 0.43884 | 0.39353
74.4 | 41.9 | 140.9 | 18 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 70.1486 | 40.8799 | 128.9945 | 278.1854 | 289.1357 | 278.3922 | 269.7486 | 275.6619 | 0.44707 | 0.22992 | 0.45035 | 0.10427 | 0.45362 | 0.49765 | 2016-09-18 08:56 | 2.2368 | 2.26 | 2016-09-18 08:57 | 2.2205 | 2.2431 | 2016-09-18 08:58 | 2.2045 | 2.2265 | 2016-09-18 08:56 | 2.2286 | 2.2515 | 2016-09-18 08:58 | 2.2125 | 2.2347 | 0.44871 | 0.12566 | 0.45198 | 0.39338
72.3 | 40.8 | 137.2 | 18 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 68.2133 | 39.8048 | 125.4689 | 278.5847 | 290.2752 | 278.6165 | 269.5717 | 275.656 | 0.46014 | 0.23025 | 0.46339 | 0.10468 | 0.46663 | 0.49805 | 2016-09-18 09:00 | 2.1732 | 2.1941 | 2016-09-18 09:01 | 2.158 | 2.1783 | 2016-09-18 09:02 | 2.143 | 2.1628 | 2016-09-18 09:00 | 2.1656 | 2.1861 | 2016-09-18 09:02 | 2.1505 | 2.1705 | 0.46177 | 0.12557 | 0.46501 | 0.39337
70.4 | 39.9 | 133.8 | 18 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 66.4634 | 38.9219 | 122.2228 | 279.3314 | 293.0309 | 278.9917 | 268.7605 | 275.4306 | 0.4731 | 0.23087 | 0.47632 | 0.10567 | 0.47953 | 0.49872 | 2016-09-18 09:04 | 2.1137 | 2.1325 | 2016-09-18 09:05 | 2.0994 | 2.1177 | 2016-09-18 09:06 | 2.0854 | 2.1032 | 2016-09-18 09:05 | 2.1066 | 2.1251 | 2016-09-18 09:06 | 2.0924 | 2.1104 | 0.47471 | 0.1252 | 0.47792 | 0.39305
68.4 | 38.7 | 130.5 | 18 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 64.6224 | 37.7395 | 119.0686 | 278.8909 | 290.9472 | 279.0648 | 269.5895 | 276.0495 | 0.48593 | 0.2305 | 0.48912 | 0.10492 | 0.4923 | 0.49885 | 2016-09-18 09:08 | 2.0579 | 2.0749 | 2016-09-18 09:09 | 2.0445 | 2.061 | 2016-09-18 09:10 | 2.0313 | 2.0474 | 2016-09-18 09:08 | 2.0512 | 2.0679 | 2016-09-18 09:10 | 2.0379 | 2.0542 | 0.48752 | 0.12559 | 0.49071 | 0.39393
66.5 | 37.6 | 127.3 | 18 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 62.8717 | 36.6508 | 116.0169 | 278.3202 | 289.183 | 278.8896 | 269.9403 | 276.2766 | 0.49863 | 0.23003 | 0.50179 | 0.10428 | 0.50493 | 0.49854 | 2016-09-18 09:12 | 2.0055 | 2.0208 | 2016-09-18 09:13 | 1.9929 | 2.0078 | 2016-09-18 09:14 | 1.9805 | 1.9951 | 2016-09-18 09:12 | 1.9992 | 2.0143 | 2016-09-18 09:14 | 1.9866 | 2.0014 | 0.50021 | 0.12575 | 0.50336 | 0.39425
56.5 | 32.5 | 109.7 | 18 | 9 | 40 | 0 | 9 | 16 | C | 83 | IZO | 53.5954 | 31.5558 | 99.4455 | 277.7218 | 293.008 | 279.1677 | 265.9076 | 275.6617 | 0.58355 | 0.22954 | 0.58644 | 0.10566 | 0.58932 | 0.49903 | 2016-09-18 09:40 | 1.7136 | 1.7216 | 2016-09-18 09:41 | 1.7052 | 1.713 | 2016-09-18 09:42 | 1.6969 | 1.7045 | 2016-09-18 09:40 | 1.7094 | 1.7173 | 2016-09-18 09:42 | 1.701 | 1.7087 | 0.585 | 0.12388 | 0.58788 | 0.39337
55.7 | 32 | 107.9 | 18 | 9 | 44 | 0 | 9 | 16 | C | 83 | IZO | 52.8482 | 31.0529 | 97.7697 | 279.8346 | 294.7679 | 279.9047 | 268.2925 | 276.1406 | 0.59506 | 0.23128 | 0.59791 | 0.1063 | 0.60075 | 0.50035 | 2016-09-18 09:44 | 1.6805 | 1.6878 | 2016-09-18 09:45 | 1.6725 | 1.6797 | 2016-09-18 09:46 | 1.6646 | 1.6716 | 2016-09-18 09:44 | 1.6765 | 1.6837 | 2016-09-18 09:46 | 1.6685 | 1.6756 | 0.59648 | 0.12499 | 0.59933 | 0.39406
54.9 | 31.5 | 105.7 | 18 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 52.0998 | 30.5493 | 95.724 | 281.6216 | 296.0483 | 279.068 | 270.4708 | 274.7691 | 0.60639 | 0.23276 | 0.6092 | 0.10676 | 0.612 | 0.49886 | 2016-09-18 09:48 | 1.6491 | 1.6558 | 2016-09-18 09:49 | 1.6415 | 1.648 | 2016-09-18 09:50 | 1.634 | 1.6404 | 2016-09-18 09:48 | 1.6453 | 1.6519 | 2016-09-18 09:50 | 1.6377 | 1.6442 | 0.6078 | 0.126 | 0.6106 | 0.3921
53.7 | 30.9 | 104 | 18 | 9 | 52 | 0 | 9 | 16 | C | 83 | IZO | 50.9743 | 29.9441 | 94.1381 | 280.2763 | 295.1158 | 279.4842 | 268.8053 | 275.5264 | 0.61756 | 0.23165 | 0.62032 | 0.10642 | 0.62307 | 0.4996 | 2016-09-18 09:52 | 1.6193 | 1.6254 | 2016-09-18 09:53 | 1.6121 | 1.6181 | 2016-09-18 09:54 | 1.6049 | 1.6108 | 2016-09-18 09:52 | 1.6157 | 1.6217 | 2016-09-18 09:54 | 1.6085 | 1.6144 | 0.61894 | 0.12523 | 0.6217 | 0.39318
52.7 | 30.3 | 102.4 | 18 | 9 | 56 | 0 | 9 | 16 | C | 83 | IZO | 50.0332 | 29.3378 | 92.6388 | 279.9205 | 293.6426 | 279.9178 | 269.3136 | 276.4421 | 0.62855 | 0.23136 | 0.63126 | 0.10589 | 0.63397 | 0.50038 | 2016-09-18 09:56 | 1.591 | 1.5966 | 2016-09-18 09:57 | 1.5841 | 1.5896 | 2016-09-18 09:58 | 1.5774 | 1.5828 | 2016-09-18 09:56 | 1.5875 | 1.5931 | 2016-09-18 09:58 | 1.5807 | 1.5862 | 0.62991 | 0.12546 | 0.63262 | 0.39448
51.8 | 29.8 | 100.6 | 18 | 10 | 0 | 0 | 9 | 16 | C | 83 | IZO | 49.1839 | 28.8317 | 90.9476 | 279.9163 | 293.437 | 279.3828 | 269.4646 | 275.8241 | 0.63935 | 0.23135 | 0.64203 | 0.10582 | 0.64469 | 0.49942 | 2016-09-18 10:00 | 1.5641 | 1.5693 | 2016-09-18 10:01 | 1.5576 | 1.5626 | 2016-09-18 10:02 | 1.5511 | 1.5561 | 2016-09-18 10:00 | 1.5608 | 1.5659 | 2016-09-18 10:02 | 1.5543 | 1.5594 | 0.64069 | 0.12554 | 0.64336 | 0.3936
51 | 29.1 | 99.3 | 18 | 10 | 4 | 0 | 9 | 16 | C | 83 | IZO | 48.4278 | 28.1217 | 89.7247 | 280.314 | 289.0914 | 280.268 | 273.5331 | 278.0314 | 0.64998 | 0.23168 | 0.6526 | 0.10425 | 0.65522 | 0.501 | 2016-09-18 10:04 | 1.5385 | 1.5433 | 2016-09-18 10:05 | 1.5323 | 1.537 | 2016-09-18 10:06 | 1.5262 | 1.5308 | 2016-09-18 10:04 | 1.5354 | 1.5401 | 2016-09-18 10:06 | 1.5293 | 1.5339 | 0.65129 | 0.12743 | 0.65391 | 0.39675
50.1 | 28.7 | 97.8 | 18 | 10 | 8 | 0 | 9 | 16 | C | 83 | IZO | 47.5756 | 27.7152 | 88.3137 | 279.6394 | 289.763 | 280.2144 | 271.8158 | 277.7946 | 0.66041 | 0.23112 | 0.66299 | 0.10449 | 0.66556 | 0.50091 | 2016-09-18 10:08 | 1.5142 | 1.5186 | 2016-09-18 10:09 | 1.5083 | 1.5126 | 2016-09-18 10:10 | 1.5025 | 1.5067 | 2016-09-18 10:08 | 1.5113 | 1.5156 | 2016-09-18 10:10 | 1.5054 | 1.5096 | 0.6617 | 0.12663 | 0.66428 | 0.39641
49.5 | 28.4 | 96.2 | 18 | 10 | 12 | 0 | 9 | 16 | C | 83 | IZO | 47.0062 | 27.4099 | 86.8084 | 280.9358 | 291.9421 | 279.5469 | 272.4285 | 276.408 | 0.67066 | 0.23219 | 0.67319 | 0.10528 | 0.67571 | 0.49971 | 2016-09-18 10:12 | 1.4911 | 1.4951 | 2016-09-18 10:13 | 1.4855 | 1.4894 | 2016-09-18 10:14 | 1.4799 | 1.4838 | 2016-09-18 10:12 | 1.4883 | 1.4923 | 2016-09-18 10:14 | 1.4827 | 1.4866 | 0.67192 | 0.12692 | 0.67445 | 0.39443
48.8 | 28.2 | 95 | 18 | 10 | 16 | 0 | 9 | 16 | C | 83 | IZO | 46.3406 | 27.2062 | 85.6778 | 281.1877 | 295.7234 | 280.0513 | 269.948 | 276.0844 | 0.68071 | 0.2324 | 0.68319 | 0.10664 | 0.68566 | 0.50061 | 2016-09-18 10:16 | 1.4691 | 1.4728 | 2016-09-18 10:17 | 1.4637 | 1.4674 | 2016-09-18 10:18 | 1.4585 | 1.462 | 2016-09-18 10:16 | 1.4664 | 1.4701 | 2016-09-18 10:18 | 1.4611 | 1.4647 | 0.68195 | 0.12576 | 0.68443 | 0.39397
48 | 27.8 | 93.8 | 18 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 45.5778 | 26.7984 | 84.5473 | 280.3553 | 295.3173 | 280.3356 | 268.7852 | 276.5434 | 0.69056 | 0.23171 | 0.69299 | 0.10649 | 0.69541 | 0.50112 | 2016-09-18 10:20 | 1.4481 | 1.4515 | 2016-09-18 10:21 | 1.443 | 1.4464 | 2016-09-18 10:22 | 1.438 | 1.4413 | 2016-09-18 10:20 | 1.4456 | 1.4489 | 2016-09-18 10:22 | 1.4405 | 1.4438 | 0.69178 | 0.12522 | 0.6942 | 0.39463
47.4 | 28.1 | 92.8 | 18 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 45.0041 | 27.1043 | 83.6065 | 280.8328 | 308.4467 | 281.1497 | 259.4693 | 274.2454 | 0.70021 | 0.23211 | 0.70259 | 0.11123 | 0.70496 | 0.50258 | 2016-09-18 10:24 | 1.4281 | 1.4313 | 2016-09-18 10:25 | 1.4233 | 1.4264 | 2016-09-18 10:26 | 1.4185 | 1.4215 | 2016-09-18 10:24 | 1.4257 | 1.4288 | 2016-09-18 10:25 | 1.4209 | 1.424 | 0.7014 | 0.12088 | 0.70378 | 0.39135
47 | 27.5 | 91.9 | 18 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 44.6207 | 26.4921 | 82.7616 | 282.7004 | 303.4528 | 282.1508 | 266.6475 | 276.7616 | 0.70966 | 0.23365 | 0.71199 | 0.10943 | 0.71431 | 0.50437 | 2016-09-18 10:28 | 1.4091 | 1.412 | 2016-09-18 10:29 | 1.4045 | 1.4074 | 2016-09-18 10:30 | 1.4 | 1.4027 | 2016-09-18 10:28 | 1.4068 | 1.4097 | 2016-09-18 10:30 | 1.4022 | 1.405 | 0.71083 | 0.12422 | 0.71315 | 0.39494
46.2 | 27.4 | 90.6 | 18 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 43.8519 | 26.3899 | 81.5455 | 281.015 | 308.1893 | 281.4651 | 259.991 | 274.7058 | 0.7189 | 0.23226 | 0.72118 | 0.11114 | 0.72344 | 0.50314 | 2016-09-18 10:32 | 1.391 | 1.3937 | 2016-09-18 10:33 | 1.3866 | 1.3892 | 2016-09-18 10:34 | 1.3823 | 1.3848 | 2016-09-18 10:32 | 1.3888 | 1.3915 | 2016-09-18 10:34 | 1.3844 | 1.387 | 0.72004 | 0.12112 | 0.72231 | 0.392

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.18 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:0.05 C:-0.13 D:0.39 Nad:-0.34 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 274.1912 | 276.6816 | 275.515 | 272.3238 | 275.1982 | All Obs | -0.12867 |  | 0.39224 |  | 0.052324 | -0.33991 | -0.52091
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
138.6 | 79.5 | 248.1 | 18 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 129.0004 | 74.6918 | 233.8714 | 272.2899 | 278.7299 | 274.6492 | 267.5696 | 273.5312 | 0.24162 | 0.22505 | 0.24509 | 0.10051 | 0.24856 | 0.49096 | 2016-09-18 07:56 | 4.1387 | 4.3291 | 2016-09-18 07:57 | 4.0801 | 4.2617 | 2016-09-18 07:58 | 4.0231 | 4.1964 | 2016-09-18 07:56 | 4.1092 | 4.2951 | 2016-09-18 07:57 | 4.0514 | 4.2288
131.1 | 74.9 | 236.3 | 18 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 121.9404 | 70.4764 | 221.7697 | 272.718 | 278.4353 | 275.3355 | 268.5136 | 274.479 | 0.25552 | 0.2254 | 0.259 | 0.10041 | 0.26248 | 0.49218 | 2016-09-18 08:00 | 3.9136 | 4.0717 | 2016-09-18 08:01 | 3.861 | 4.0122 | 2016-09-18 08:02 | 3.8098 | 3.9544 | 2016-09-18 08:01 | 3.8871 | 4.0417 | 2016-09-18 08:02 | 3.8352 | 3.9831
124.8 | 71 | 225.3 | 18 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 116.0587 | 66.9113 | 210.5236 | 274.5648 | 279.6696 | 275.4461 | 270.8003 | 274.3164 | 0.26945 | 0.22693 | 0.27293 | 0.10085 | 0.27642 | 0.49238 | 2016-09-18 08:04 | 3.7113 | 3.8438 | 2016-09-18 08:05 | 3.6639 | 3.7909 | 2016-09-18 08:06 | 3.6177 | 3.7394 | 2016-09-18 08:04 | 3.6874 | 3.8171 | 2016-09-18 08:06 | 3.6406 | 3.765
118.5 | 67.3 | 215.5 | 18 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 110.1991 | 63.5393 | 200.7217 | 274.4481 | 279.5475 | 276.1206 | 270.6602 | 275.2009 | 0.28339 | 0.22683 | 0.28688 | 0.10081 | 0.29036 | 0.49359 | 2016-09-18 08:08 | 3.5287 | 3.6407 | 2016-09-18 08:09 | 3.4858 | 3.5934 | 2016-09-18 08:10 | 3.444 | 3.5473 | 2016-09-18 08:08 | 3.5071 | 3.6169 | 2016-09-18 08:10 | 3.4648 | 3.5702
112.6 | 64.2 | 206.1 | 18 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.7328 | 60.7026 | 191.3785 | 273.7043 | 281.1459 | 276.0081 | 268.0811 | 274.6631 | 0.29733 | 0.22622 | 0.30081 | 0.10138 | 0.30429 | 0.49339 | 2016-09-18 08:12 | 3.3633 | 3.4587 | 2016-09-18 08:13 | 3.3244 | 3.4162 | 2016-09-18 08:14 | 3.2863 | 3.3747 | 2016-09-18 08:12 | 3.3437 | 3.4373 | 2016-09-18 08:13 | 3.3052 | 3.3953
107.5 | 61.2 | 197.4 | 18 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 100.0078 | 57.945 | 182.8159 | 273.8372 | 281.1074 | 275.7814 | 268.3284 | 274.3948 | 0.31125 | 0.22633 | 0.31472 | 0.10137 | 0.31819 | 0.49298 | 2016-09-18 08:16 | 3.2129 | 3.2948 | 2016-09-18 08:17 | 3.1774 | 3.2563 | 2016-09-18 08:18 | 3.1427 | 3.2188 | 2016-09-18 08:17 | 3.1951 | 3.2754 | 2016-09-18 08:18 | 3.16 | 3.2374
102.7 | 58.3 | 189.4 | 18 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 95.5525 | 55.2647 | 174.9317 | 273.3634 | 279.5099 | 275.4287 | 268.7093 | 274.3614 | 0.32513 | 0.22594 | 0.3286 | 0.10079 | 0.33206 | 0.49235 | 2016-09-18 08:20 | 3.0757 | 3.1464 | 2016-09-18 08:21 | 3.0432 | 3.1114 | 2016-09-18 08:22 | 3.0115 | 3.0773 | 2016-09-18 08:20 | 3.0594 | 3.1288 | 2016-09-18 08:22 | 3.0273 | 3.0943
98.5 | 55.7 | 182.4 | 18 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 91.6466 | 52.8504 | 168.0292 | 273.5647 | 278.2395 | 275.6922 | 270.0372 | 275.0165 | 0.33897 | 0.2261 | 0.34243 | 0.10034 | 0.34587 | 0.49282 | 2016-09-18 08:24 | 2.9501 | 3.0115 | 2016-09-18 08:25 | 2.9203 | 2.9796 | 2016-09-18 08:26 | 2.8912 | 2.9485 | 2016-09-18 08:24 | 2.9351 | 2.9955 | 2016-09-18 08:26 | 2.9057 | 2.964
94.8 | 53.6 | 175.7 | 18 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 88.2152 | 50.889 | 161.4548 | 274.399 | 279.4549 | 275.4838 | 270.5661 | 274.4516 | 0.35276 | 0.22679 | 0.3562 | 0.10077 | 0.35963 | 0.49245 | 2016-09-18 08:28 | 2.8348 | 2.8883 | 2016-09-18 08:29 | 2.8074 | 2.8592 | 2016-09-18 08:30 | 2.7806 | 2.8308 | 2016-09-18 08:28 | 2.821 | 2.8737 | 2016-09-18 08:29 | 2.794 | 2.8449
91.1 | 51.6 | 170 | 18 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 84.7873 | 49.0071 | 155.8859 | 274.0159 | 279.8513 | 276.2948 | 269.5706 | 275.3703 | 0.36649 | 0.22647 | 0.36991 | 0.10092 | 0.37332 | 0.4939 | 2016-09-18 08:32 | 2.7286 | 2.7756 | 2016-09-18 08:33 | 2.7034 | 2.7489 | 2016-09-18 08:34 | 2.6787 | 2.7227 | 2016-09-18 08:33 | 2.7159 | 2.7622 | 2016-09-18 08:34 | 2.691 | 2.7357
87.9 | 49.7 | 164.1 | 18 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 81.8189 | 47.2081 | 150.1131 | 274.5178 | 279.5617 | 275.7389 | 270.6775 | 274.7495 | 0.38014 | 0.22689 | 0.38354 | 0.10081 | 0.38694 | 0.4929 | 2016-09-18 08:36 | 2.6306 | 2.672 | 2016-09-18 08:37 | 2.6073 | 2.6474 | 2016-09-18 08:38 | 2.5844 | 2.6233 | 2016-09-18 08:36 | 2.6189 | 2.6596 | 2016-09-18 08:38 | 2.5958 | 2.6353
85 | 47.8 | 159 | 18 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 79.1343 | 45.398 | 145.1413 | 275.2987 | 277.6521 | 276.0235 | 273.5345 | 275.5908 | 0.39371 | 0.22754 | 0.39709 | 0.10012 | 0.40047 | 0.49341 | 2016-09-18 08:40 | 2.5399 | 2.5765 | 2016-09-18 08:41 | 2.5183 | 2.5538 | 2016-09-18 08:42 | 2.4971 | 2.5316 | 2016-09-18 08:40 | 2.5291 | 2.5651 | 2016-09-18 08:42 | 2.5077 | 2.5426
82.1 | 46.3 | 154.1 | 18 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 76.4628 | 43.9605 | 140.3992 | 275.1677 | 278.4872 | 276.0102 | 272.6501 | 275.3648 | 0.4072 | 0.22743 | 0.41056 | 0.10043 | 0.41391 | 0.49339 | 2016-09-18 08:44 | 2.4558 | 2.4883 | 2016-09-18 08:45 | 2.4357 | 2.4673 | 2016-09-18 08:46 | 2.416 | 2.4467 | 2016-09-18 08:44 | 2.4457 | 2.4777 | 2016-09-18 08:46 | 2.4258 | 2.4569
79.4 | 44.7 | 149.4 | 18 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 73.98 | 42.4192 | 135.842 | 275.0286 | 276.8898 | 275.6463 | 273.6348 | 275.3142 | 0.42059 | 0.22731 | 0.42392 | 0.099849 | 0.42725 | 0.49274 | 2016-09-18 08:48 | 2.3776 | 2.4066 | 2016-09-18 08:49 | 2.3589 | 2.387 | 2016-09-18 08:50 | 2.3406 | 2.3679 | 2016-09-18 08:49 | 2.3682 | 2.3968 | 2016-09-18 08:50 | 2.3497 | 2.3774
76.7 | 43.3 | 145.1 | 18 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 71.4933 | 41.0639 | 131.7152 | 273.9674 | 276.2668 | 275.582 | 272.2304 | 275.3925 | 0.43388 | 0.22643 | 0.43719 | 0.099625 | 0.44049 | 0.49262 | 2016-09-18 08:52 | 2.3048 | 2.3306 | 2016-09-18 08:53 | 2.2873 | 2.3125 | 2016-09-18 08:54 | 2.2702 | 2.2947 | 2016-09-18 08:52 | 2.296 | 2.3215 | 2016-09-18 08:54 | 2.2787 | 2.3035
74.4 | 41.9 | 140.9 | 18 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 69.3737 | 39.7019 | 127.7055 | 273.9941 | 274.4253 | 275.1214 | 273.6994 | 275.282 | 0.44707 | 0.22646 | 0.45035 | 0.098961 | 0.45362 | 0.4918 | 2016-09-18 08:56 | 2.2368 | 2.26 | 2016-09-18 08:57 | 2.2205 | 2.2431 | 2016-09-18 08:58 | 2.2045 | 2.2265 | 2016-09-18 08:56 | 2.2286 | 2.2515 | 2016-09-18 08:58 | 2.2125 | 2.2347
72.3 | 40.8 | 137.2 | 18 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 67.4384 | 38.6269 | 124.1799 | 274.2708 | 275.1388 | 275.2518 | 273.6348 | 275.2662 | 0.46014 | 0.22669 | 0.46339 | 0.099218 | 0.46663 | 0.49203 | 2016-09-18 09:00 | 2.1732 | 2.1941 | 2016-09-18 09:01 | 2.158 | 2.1783 | 2016-09-18 09:02 | 2.143 | 2.1628 | 2016-09-18 09:00 | 2.1656 | 2.1861 | 2016-09-18 09:02 | 2.1505 | 2.1705
70.4 | 39.9 | 133.8 | 18 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 65.6886 | 37.744 | 120.9339 | 274.896 | 277.4723 | 275.5341 | 272.9349 | 275.031 | 0.4731 | 0.2272 | 0.47632 | 0.10006 | 0.47953 | 0.49254 | 2016-09-18 09:04 | 2.1137 | 2.1325 | 2016-09-18 09:05 | 2.0994 | 2.1177 | 2016-09-18 09:06 | 2.0854 | 2.1032 | 2016-09-18 09:05 | 2.1066 | 2.1251 | 2016-09-18 09:06 | 2.0924 | 2.1104
68.4 | 38.7 | 130.5 | 18 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 63.8476 | 36.5616 | 117.7796 | 274.3353 | 274.9704 | 275.5151 | 273.8742 | 275.6403 | 0.48593 | 0.22674 | 0.48912 | 0.099157 | 0.4923 | 0.4925 | 2016-09-18 09:08 | 2.0579 | 2.0749 | 2016-09-18 09:09 | 2.0445 | 2.061 | 2016-09-18 09:10 | 2.0313 | 2.0474 | 2016-09-18 09:08 | 2.0512 | 2.0679 | 2016-09-18 09:10 | 2.0379 | 2.0542
66.5 | 37.6 | 127.3 | 18 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 62.0969 | 35.4728 | 114.728 | 273.6455 | 272.7924 | 275.2488 | 274.334 | 275.8578 | 0.49863 | 0.22617 | 0.50179 | 0.098372 | 0.50493 | 0.49203 | 2016-09-18 09:12 | 2.0055 | 2.0208 | 2016-09-18 09:13 | 1.9929 | 2.0078 | 2016-09-18 09:14 | 1.9805 | 1.9951 | 2016-09-18 09:12 | 1.9992 | 2.0143 | 2016-09-18 09:14 | 1.9866 | 2.0014
56.5 | 32.5 | 109.7 | 18 | 9 | 40 | 0 | 9 | 16 | C | 83 | IZO | 52.8205 | 30.3778 | 98.1565 | 272.2509 | 273.8521 | 274.9184 | 271.0294 | 275.1795 | 0.58355 | 0.22502 | 0.58644 | 0.098754 | 0.58932 | 0.49144 | 2016-09-18 09:40 | 1.7136 | 1.7216 | 2016-09-18 09:41 | 1.7052 | 1.713 | 2016-09-18 09:42 | 1.6969 | 1.7045 | 2016-09-18 09:40 | 1.7094 | 1.7173 | 2016-09-18 09:42 | 1.701 | 1.7087
55.7 | 32 | 107.9 | 18 | 9 | 44 | 0 | 9 | 16 | C | 83 | IZO | 52.0734 | 29.8749 | 96.4808 | 274.2559 | 275.2376 | 275.5731 | 273.5128 | 275.6498 | 0.59506 | 0.22667 | 0.59791 | 0.099254 | 0.60075 | 0.49261 | 2016-09-18 09:44 | 1.6805 | 1.6878 | 2016-09-18 09:45 | 1.6725 | 1.6797 | 2016-09-18 09:46 | 1.6646 | 1.6716 | 2016-09-18 09:44 | 1.6765 | 1.6837 | 2016-09-18 09:46 | 1.6685 | 1.6756
54.9 | 31.5 | 105.7 | 18 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 51.325 | 29.3714 | 94.435 | 275.9366 | 276.149 | 274.6552 | 275.7881 | 274.2699 | 0.60639 | 0.22806 | 0.6092 | 0.099582 | 0.612 | 0.49097 | 2016-09-18 09:48 | 1.6491 | 1.6558 | 2016-09-18 09:49 | 1.6415 | 1.648 | 2016-09-18 09:50 | 1.634 | 1.6404 | 2016-09-18 09:48 | 1.6453 | 1.6519 | 2016-09-18 09:50 | 1.6377 | 1.6442
53.7 | 30.9 | 104 | 18 | 9 | 52 | 0 | 9 | 16 | C | 83 | IZO | 50.1995 | 28.7662 | 92.8491 | 274.4867 | 274.8532 | 274.9915 | 274.2181 | 275.019 | 0.61756 | 0.22686 | 0.62032 | 0.099115 | 0.62307 | 0.49157 | 2016-09-18 09:52 | 1.6193 | 1.6254 | 2016-09-18 09:53 | 1.6121 | 1.6181 | 2016-09-18 09:54 | 1.6049 | 1.6108 | 2016-09-18 09:52 | 1.6157 | 1.6217 | 2016-09-18 09:54 | 1.6085 | 1.6144
52.7 | 30.3 | 102.4 | 18 | 9 | 56 | 0 | 9 | 16 | C | 83 | IZO | 49.2583 | 28.1599 | 91.3499 | 274.0279 | 273.0227 | 275.3466 | 274.8203 | 275.9266 | 0.62855 | 0.22648 | 0.63126 | 0.098455 | 0.63397 | 0.4922 | 2016-09-18 09:56 | 1.591 | 1.5966 | 2016-09-18 09:57 | 1.5841 | 1.5896 | 2016-09-18 09:58 | 1.5774 | 1.5828 | 2016-09-18 09:56 | 1.5875 | 1.5931 | 2016-09-18 09:58 | 1.5807 | 1.5862
51.8 | 29.8 | 100.6 | 18 | 10 | 0 | 0 | 9 | 16 | C | 83 | IZO | 48.409 | 27.6537 | 89.6587 | 273.9224 | 272.4655 | 274.7343 | 275.0636 | 275.3006 | 0.63935 | 0.2264 | 0.64203 | 0.098254 | 0.64469 | 0.49111 | 2016-09-18 10:00 | 1.5641 | 1.5693 | 2016-09-18 10:01 | 1.5576 | 1.5626 | 2016-09-18 10:02 | 1.5511 | 1.5561 | 2016-09-18 10:00 | 1.5608 | 1.5659 | 2016-09-18 10:02 | 1.5543 | 1.5594
51 | 29.1 | 99.3 | 18 | 10 | 4 | 0 | 9 | 16 | C | 83 | IZO | 47.6529 | 26.9438 | 88.4358 | 274.2204 | 267.7744 | 275.5436 | 279.2229 | 277.5001 | 0.64998 | 0.22664 | 0.6526 | 0.096562 | 0.65522 | 0.49256 | 2016-09-18 10:04 | 1.5385 | 1.5433 | 2016-09-18 10:05 | 1.5323 | 1.537 | 2016-09-18 10:06 | 1.5262 | 1.5308 | 2016-09-18 10:04 | 1.5354 | 1.5401 | 2016-09-18 10:06 | 1.5293 | 1.5339
50.1 | 28.7 | 97.8 | 18 | 10 | 8 | 0 | 9 | 16 | C | 83 | IZO | 46.8007 | 26.5373 | 87.0248 | 273.4479 | 268.1067 | 275.4154 | 277.5947 | 277.2557 | 0.66041 | 0.22601 | 0.66299 | 0.096682 | 0.66556 | 0.49233 | 2016-09-18 10:08 | 1.5142 | 1.5186 | 2016-09-18 10:09 | 1.5083 | 1.5126 | 2016-09-18 10:10 | 1.5025 | 1.5067 | 2016-09-18 10:08 | 1.5113 | 1.5156 | 2016-09-18 10:10 | 1.5054 | 1.5096
49.5 | 28.4 | 96.2 | 18 | 10 | 12 | 0 | 9 | 16 | C | 83 | IZO | 46.2314 | 26.232 | 85.5195 | 274.6483 | 269.9527 | 274.6747 | 278.2948 | 275.8615 | 0.67066 | 0.227 | 0.67319 | 0.097348 | 0.67571 | 0.491 | 2016-09-18 10:12 | 1.4911 | 1.4951 | 2016-09-18 10:13 | 1.4855 | 1.4894 | 2016-09-18 10:14 | 1.4799 | 1.4838 | 2016-09-18 10:12 | 1.4883 | 1.4923 | 2016-09-18 10:14 | 1.4827 | 1.4866
48.8 | 28.2 | 95 | 18 | 10 | 16 | 0 | 9 | 16 | C | 83 | IZO | 45.5657 | 26.0283 | 84.3888 | 274.806 | 273.4073 | 275.1074 | 275.9 | 275.5307 | 0.68071 | 0.22713 | 0.68319 | 0.098594 | 0.68566 | 0.49178 | 2016-09-18 10:16 | 1.4691 | 1.4728 | 2016-09-18 10:17 | 1.4637 | 1.4674 | 2016-09-18 10:18 | 1.4585 | 1.462 | 2016-09-18 10:16 | 1.4664 | 1.4701 | 2016-09-18 10:18 | 1.4611 | 1.4647
48 | 27.8 | 93.8 | 18 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 44.803 | 25.6205 | 83.2584 | 273.8812 | 272.681 | 275.3214 | 274.8212 | 275.9825 | 0.69056 | 0.22636 | 0.69299 | 0.098332 | 0.69541 | 0.49216 | 2016-09-18 10:20 | 1.4481 | 1.4515 | 2016-09-18 10:21 | 1.443 | 1.4464 | 2016-09-18 10:22 | 1.438 | 1.4413 | 2016-09-18 10:20 | 1.4456 | 1.4489 | 2016-09-18 10:22 | 1.4405 | 1.4438
47.4 | 28.1 | 92.8 | 18 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 44.2292 | 25.9264 | 82.3176 | 274.2682 | 285.4968 | 276.0666 | 265.5875 | 273.6774 | 0.70021 | 0.22668 | 0.70259 | 0.10295 | 0.70496 | 0.49349 | 2016-09-18 10:24 | 1.4281 | 1.4313 | 2016-09-18 10:25 | 1.4233 | 1.4264 | 2016-09-18 10:26 | 1.4185 | 1.4215 | 2016-09-18 10:24 | 1.4257 | 1.4288 | 2016-09-18 10:25 | 1.4209 | 1.424
47 | 27.5 | 91.9 | 18 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 43.8459 | 25.3142 | 81.4727 | 276.0473 | 280.196 | 277.0003 | 272.8461 | 276.1868 | 0.70966 | 0.22815 | 0.71199 | 0.10104 | 0.71431 | 0.49516 | 2016-09-18 10:28 | 1.4091 | 1.412 | 2016-09-18 10:29 | 1.4045 | 1.4074 | 2016-09-18 10:30 | 1.4 | 1.4027 | 2016-09-18 10:28 | 1.4068 | 1.4097 | 2016-09-18 10:30 | 1.4022 | 1.405
46.2 | 27.4 | 90.6 | 18 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 43.077 | 25.212 | 80.2565 | 274.2752 | 284.6324 | 276.2487 | 266.2682 | 274.1243 | 0.7189 | 0.22669 | 0.72118 | 0.10264 | 0.72344 | 0.49382 | 2016-09-18 10:32 | 1.391 | 1.3937 | 2016-09-18 10:33 | 1.3866 | 1.3892 | 2016-09-18 10:34 | 1.3823 | 1.3848 | 2016-09-18 10:32 | 1.3888 | 1.3915 | 2016-09-18 10:34 | 1.3844 | 1.387

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.7876 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.8265 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.36061 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.427 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46588 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## comp page
Komyhr adjusted |  |  |  |  | IUP MULTIPLY |  |  |  |  | B&P B MULTIPLY |  |  |  |  | D&M MULTIPLY
276.5924 | 282.244 | 276.2315 | 272.0444 | 274.6431 | 278.7668 | 292.7235 | 279.0781 | 268.0241 | 275.6078 | 277.0634 | 287.2712 | 277.0905 | 269.0313 | 274.4785 | 280.9471 | 295.1725 | 279.3321 | 269.9925 | 275.3315
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
272.4136 | 276.4705 | 273.6218 | 269.3723 | 272.7951 | 274.5551 | 286.7356 | 276.4415 | 265.3915 | 273.7532 | 272.8775 | 281.3949 | 274.4727 | 266.3888 | 272.6315 | 276.7025 | 289.1346 | 276.6931 | 267.3405 | 273.4788
272.9676 | 276.6245 | 274.4004 | 270.2082 | 273.7504 | 275.1135 | 286.8954 | 277.2281 | 266.2149 | 274.7119 | 273.4324 | 281.5517 | 275.2537 | 267.2154 | 273.5863 | 277.2652 | 289.2957 | 277.4804 | 268.1701 | 274.4365
274.9296 | 278.2535 | 274.6093 | 272.4067 | 273.5992 | 277.0909 | 288.5848 | 277.4392 | 268.381 | 274.5602 | 275.3978 | 283.2097 | 275.4633 | 269.3895 | 273.4352 | 279.2581 | 290.9993 | 277.6917 | 270.352 | 274.285
274.9435 | 278.5749 | 275.3764 | 272.142 | 274.4914 | 277.1049 | 288.9182 | 278.2142 | 268.1202 | 275.4555 | 275.4117 | 283.5368 | 276.2327 | 269.1278 | 274.3268 | 279.2722 | 291.3354 | 278.4674 | 270.0894 | 275.1794
274.3352 | 280.555 | 275.3645 | 269.402 | 273.9663 | 276.4918 | 290.9718 | 278.2022 | 265.4207 | 274.9286 | 274.8023 | 285.5521 | 276.2208 | 266.4182 | 273.802 | 278.6543 | 293.4062 | 278.4554 | 267.37 | 274.6529
274.5965 | 280.9559 | 275.2394 | 269.531 | 273.7096 | 276.7551 | 291.3876 | 278.0758 | 265.5477 | 274.6709 | 275.0641 | 285.9603 | 276.0953 | 266.5456 | 273.5455 | 278.9197 | 293.8255 | 278.3288 | 267.498 | 274.3956
274.2555 | 279.8526 | 274.9892 | 269.7959 | 273.687 | 276.4115 | 290.2433 | 277.823 | 265.8088 | 274.6483 | 274.7225 | 284.8373 | 275.8443 | 266.8076 | 273.5229 | 278.5734 | 292.6716 | 278.0759 | 267.7609 | 274.373
274.584 | 279.0633 | 275.3486 | 271.0225 | 274.3504 | 276.7426 | 289.4247 | 278.1861 | 267.0172 | 275.314 | 275.0516 | 284.0339 | 276.2049 | 268.0206 | 274.1859 | 278.9071 | 291.8462 | 278.4393 | 268.9782 | 275.038
275.5401 | 280.6689 | 275.2405 | 271.4386 | 273.7981 | 277.7062 | 291.0899 | 278.0769 | 267.4272 | 274.7598 | 276.0093 | 285.6681 | 276.0964 | 268.4322 | 273.6339 | 279.8782 | 293.5253 | 278.33 | 269.3913 | 274.4843
275.2876 | 281.4829 | 276.141 | 270.3082 | 274.724 | 277.4517 | 291.9341 | 278.9866 | 266.3135 | 275.6889 | 275.7564 | 286.4966 | 276.9997 | 267.3142 | 274.5593 | 279.6218 | 294.3766 | 279.2405 | 268.2693 | 275.4126
275.9125 | 281.633 | 275.6878 | 271.3122 | 274.1158 | 278.0816 | 292.0898 | 278.5288 | 267.3027 | 275.0785 | 276.3824 | 286.6494 | 276.5451 | 268.3072 | 273.9514 | 280.2565 | 294.5336 | 278.7823 | 269.2658 | 274.8028
276.8137 | 280.2187 | 276.0662 | 274.0935 | 274.9644 | 278.9898 | 290.623 | 278.9111 | 270.0428 | 275.9302 | 277.2851 | 285.2098 | 276.9247 | 271.0576 | 274.7995 | 281.1719 | 293.0545 | 279.1649 | 272.0261 | 275.6535
276.8091 | 281.4479 | 276.1489 | 273.0779 | 274.7495 | 278.9852 | 291.8978 | 278.9947 | 269.0423 | 275.7145 | 277.2805 | 286.461 | 277.0076 | 270.0533 | 274.5847 | 281.1672 | 294.34 | 279.2486 | 271.0182 | 275.4381
276.7957 | 280.3287 | 275.884 | 273.9605 | 274.7092 | 278.9716 | 290.7371 | 278.727 | 269.9118 | 275.674 | 277.267 | 285.3218 | 276.7419 | 270.9261 | 274.5445 | 281.1535 | 293.1695 | 278.9806 | 271.8941 | 275.3977
275.8664 | 280.1458 | 275.9148 | 272.4189 | 274.7973 | 278.0351 | 290.5474 | 278.7582 | 268.393 | 275.7624 | 276.3362 | 285.1357 | 276.7728 | 269.4016 | 274.6325 | 280.2097 | 292.9782 | 279.0119 | 270.3641 | 275.486
276.0155 | 278.7846 | 275.5526 | 273.7948 | 274.6971 | 278.1854 | 289.1357 | 278.3922 | 269.7486 | 275.6619 | 276.4856 | 283.7502 | 276.4094 | 270.7623 | 274.5324 | 280.3611 | 291.5547 | 278.6455 | 271.7297 | 275.3855
276.4117 | 279.8834 | 275.7746 | 273.6153 | 274.6912 | 278.5847 | 290.2752 | 278.6165 | 269.5717 | 275.656 | 276.8824 | 284.8686 | 276.6322 | 270.5847 | 274.5264 | 280.7636 | 292.7038 | 278.8701 | 271.5515 | 275.3796
277.1526 | 282.5404 | 276.146 | 272.7919 | 274.4666 | 279.3314 | 293.0309 | 278.9917 | 268.7605 | 275.4306 | 277.6245 | 287.5729 | 277.0047 | 269.7705 | 274.302 | 281.5161 | 295.4825 | 279.2457 | 270.7343 | 275.1545
276.7156 | 280.5312 | 276.2183 | 273.6334 | 275.0833 | 278.8909 | 290.9472 | 279.0648 | 269.5895 | 276.0495 | 277.1868 | 285.528 | 277.0773 | 270.6026 | 274.9184 | 281.0722 | 293.3813 | 279.3188 | 271.5694 | 275.7728
276.1493 | 278.8303 | 276.0449 | 273.9894 | 275.3096 | 278.3202 | 289.183 | 278.8896 | 269.9403 | 276.2766 | 276.6196 | 283.7967 | 276.9033 | 270.9547 | 275.1445 | 280.497 | 291.6024 | 279.1434 | 271.9228 | 275.9996
275.5555 | 282.5183 | 276.3202 | 269.8962 | 274.6969 | 277.7218 | 293.008 | 279.1677 | 265.9076 | 275.6617 | 276.0248 | 287.5504 | 277.1794 | 266.9068 | 274.5322 | 279.8939 | 295.4594 | 279.4218 | 267.8604 | 275.3854
277.6519 | 284.2152 | 277.0497 | 272.3169 | 275.1741 | 279.8346 | 294.7679 | 279.9047 | 268.2925 | 276.1406 | 278.1247 | 289.2776 | 277.9112 | 269.3008 | 275.0091 | 282.0232 | 297.2341 | 280.1595 | 270.2629 | 275.8638
279.4249 | 285.4498 | 276.2215 | 274.5279 | 273.8074 | 281.6216 | 296.0483 | 279.068 | 270.4708 | 274.7691 | 279.9008 | 290.5341 | 277.0804 | 271.4872 | 273.6432 | 283.8242 | 298.5252 | 279.322 | 272.4572 | 274.4937
278.0902 | 284.5506 | 276.6334 | 272.8374 | 274.562 | 280.2763 | 295.1158 | 279.4842 | 268.8053 | 275.5264 | 278.5637 | 289.619 | 277.4937 | 269.8155 | 274.3974 | 282.4684 | 297.5848 | 279.7385 | 270.7795 | 275.2502
277.7372 | 283.1302 | 277.0626 | 273.3533 | 275.4746 | 279.9205 | 293.6426 | 279.9178 | 269.3136 | 276.4421 | 278.2101 | 288.1733 | 277.9242 | 270.3257 | 275.3094 | 282.1099 | 296.0994 | 280.1726 | 271.2915 | 276.165
277.733 | 282.932 | 276.5331 | 273.5066 | 274.8587 | 279.9163 | 293.437 | 279.3828 | 269.4646 | 275.8241 | 278.2059 | 287.9715 | 277.393 | 270.4772 | 274.6939 | 282.1056 | 295.8921 | 279.637 | 271.4436 | 275.5476
278.1276 | 278.7419 | 277.4093 | 277.6361 | 277.0583 | 280.314 | 289.0914 | 280.268 | 273.5331 | 278.0314 | 278.6012 | 283.7068 | 278.272 | 274.561 | 276.8922 | 282.5064 | 291.5101 | 280.5231 | 275.542 | 277.7527
277.4582 | 279.3895 | 277.3562 | 275.8931 | 276.8224 | 279.6394 | 289.763 | 280.2144 | 271.8158 | 277.7946 | 277.9307 | 284.3659 | 278.2186 | 272.8373 | 276.6564 | 281.8265 | 292.1873 | 280.4694 | 273.8121 | 277.5161
278.7445 | 281.4905 | 276.6955 | 276.515 | 275.4405 | 280.9358 | 291.9421 | 279.5469 | 272.4285 | 276.408 | 279.2191 | 286.5044 | 277.5559 | 273.4523 | 275.2754 | 283.133 | 294.3846 | 279.8013 | 274.4293 | 276.1309
278.9944 | 285.1365 | 277.1947 | 273.9973 | 275.1181 | 281.1877 | 295.7234 | 280.0513 | 269.948 | 276.0844 | 279.4695 | 290.2153 | 278.0567 | 270.9625 | 274.9532 | 283.3869 | 298.1976 | 280.3061 | 271.9306 | 275.8077
278.1685 | 284.7449 | 277.4762 | 272.817 | 275.5755 | 280.3553 | 295.3173 | 280.3356 | 268.7852 | 276.5434 | 278.6422 | 289.8167 | 278.3391 | 269.7953 | 275.4103 | 282.548 | 297.788 | 280.5908 | 270.7592 | 276.2662
278.6423 | 297.4043 | 278.2819 | 263.3613 | 273.2855 | 280.8328 | 308.4467 | 281.1497 | 259.4693 | 274.2454 | 279.1168 | 302.7016 | 279.1473 | 260.4444 | 273.1217 | 283.0292 | 311.0273 | 281.4055 | 261.3749 | 273.9705
280.4953 | 292.5892 | 279.2728 | 270.6472 | 275.793 | 282.7004 | 303.4528 | 282.1508 | 266.6475 | 276.7616 | 280.973 | 297.8007 | 280.1413 | 267.6496 | 275.6276 | 284.9114 | 305.9916 | 282.4076 | 268.6058 | 276.4842
278.8231 | 297.1562 | 278.5941 | 263.8909 | 273.7443 | 281.015 | 308.1893 | 281.4651 | 259.991 | 274.7058 | 279.2979 | 302.449 | 279.4605 | 260.968 | 273.5802 | 283.2129 | 310.7678 | 281.7212 | 261.9004 | 274.4304